## Example
SUMMARY
| Assay |  |  |  |  | Units |  |  | APS Absolute
| Soluble FMS Like Tyrosine Kinase 1 |  |  |  |  | ng/L |  |  | 10
| Matrix | [7 empty] | APS Relative
| Serum | [7 empty] | 10%
| Instrument/Device | [7 empty] | APS Cutoff
| Some manufacturer device | [7 empty] | 100
Method principle
Chemiluminescence
IFU Insert Version
Evaluation Period
12/11/2023
Evaluation Completed By
Comments and Assessment
Decision
| Approved By | [6 empty] | Approval Date
Commencement Date
COMARISON USING PATIENT SAMPLES
Reference Method Instrument/Device
Reference Method Principle
Reference Method IFU Version
Diagnostic Cutoff for Concordance Assessment (Enter up to 4)
|  |  | Reference Thresholds |  |  | Proposed Thresholds
|  |  | 3030 |  |  | 3020
|  |  | Reference Method |  |  | Proposed Method
Reagent Lot No.
Calibrator Lot No.
|  |  | Reference Method |  |  | Proposed Method | [6 empty] | Reference Method |  |  | Proposed Method |  |  |  |  | Concordance |  |  |  |  | Concordance
| Sample ID |  | Operator | 1st rep | 2nd rep | Operator | 1st rep | 2nd rep | Inclusion | Comment |  |  | X R1 | X R2 | Mean X | Y R1 | Y R2 | Mean Y |  |  | Cat x | Cat y | n x n |  |  | x\y | 10 | 20 | 30 | 40 | 50 | Total
| 2229705716 |  |  | 10583 | 10172 | JS | 10737 | 10746 | INCLUDE |  |  |  | 10583 | 10172 | 10377.5 | 10737 | 10746 | 10741.5 |  |  | 2 | 20 | 22 |  |  | 1 | 18 | 0 | 0 | 0 | 0 | 18
| 2229706564 |  |  | 3879 | 4233 | JS | 3934 | 3763 | INCLUDE |  |  |  | 3879 | 4233 | 4056 | 3934 | 3763 | 3848.5 |  |  | 2 | 20 | 22 |  |  | 2 | 0 | 21 | 0 | 0 | 0 | 21
| 2229804242 |  |  | 2674 | 2634 | JS | 2679 | 2632 | INCLUDE |  |  |  | 2674 | 2634 | 2654 | 2679 | 2632 | 2655.5 |  |  | 1 | 10 | 11 |  |  | 3 | 0 | 0 | 0 | 0 | 0 | 0
| 2229801511 |  |  | 4865 | 4629 | JS | 4932 | 5449 | INCLUDE |  |  |  | 4865 | 4629 | 4747 | 4932 | 5449 | 5190.5 |  |  | 2 | 20 | 22 |  |  | 4 | 0 | 0 | 0 | 0 | 0 | 0
| 2229805263 |  |  | 1471 | 1448 | JS | 1546 | 1563 | INCLUDE |  |  |  | 1471 | 1448 | 1459.5 | 1546 | 1563 | 1554.5 |  |  | 1 | 10 | 11 |  |  | 5 | 0 | 0 | 0 | 0 | 0 | 0
| 2229805226 |  |  | 6018 | 5742 | JS | 5985 | 6065 | INCLUDE |  |  |  | 6018 | 5742 | 5880 | 5985 | 6065 | 6025 |  |  | 2 | 20 | 22 |  |  | Total | 18 | 21 | 0 | 0 | 0 | 39
| 2229804358 |  |  | 3891 | 3851 | JS | 3922 | 3801 | INCLUDE |  |  |  | 3891 | 3851 | 3871 | 3922 | 3801 | 3861.5 |  |  | 2 | 20 | 22
| 2229803634 |  |  | 2519 | 2403 | JS | 2593 | 2474 | INCLUDE |  |  |  | 2519 | 2403 | 2461 | 2593 | 2474 | 2533.5 |  |  | 1 | 10 | 11
| 2229804436 |  |  | 1867 | 1737 | JS | 1880 | 1836 | INCLUDE |  |  |  | 1867 | 1737 | 1802 | 1880 | 1836 | 1858 |  |  | 1 | 10 | 11 |  |  | Agreement | 18 | 21 | 0 | 0 | 0 | 39
| 2229804060 |  |  | 1565 | 1650 | JS | 1622 | 1650 | INCLUDE |  |  |  | 1565 | 1650 | 1607.5 | 1622 | 1650 | 1636 |  |  | 1 | 10 | 11 |  |  | Agreement by chance | 8.3077 | 11.3077 | 0 | 0 | 0 | 19.6154
| 2229803948 |  |  | 4267 | 4391 | JS | 4282 | 4401 | INCLUDE |  |  |  | 4267 | 4391 | 4329 | 4282 | 4401 | 4341.5 |  |  | 2 | 20 | 22
| 2229802716 |  |  | 7862 | 7961 | JS | 8076 | 8294 | INCLUDE |  |  |  | 7862 | 7961 | 7911.5 | 8076 | 8294 | 8185 |  |  | 2 | 20 | 22
| 2229806103 |  |  | 1831 | 1766 | JS | 1813 | 1694 | INCLUDE |  |  |  | 1831 | 1766 | 1798.5 | 1813 | 1694 | 1753.5 |  |  | 1 | 10 | 11 |  |  | Proportion in agreement | 1
| 2229807634 |  |  | 6651 | 6357 | JS | 6805 | 6781 | INCLUDE |  |  |  | 6651 | 6357 | 6504 | 6805 | 6781 | 6793 |  |  | 2 | 20 | 22 |  |  | Proportion by chance | 0.50296
| 2229903100 |  |  | 9130 | 9883 | JS | 9275 | 8837 | INCLUDE |  |  |  | 9130 | 9883 | 9506.5 | 9275 | 8837 | 9056 |  |  | 2 | 20 | 22 |  |  | Cohens Kappa | 1
| 2229903874 |  |  | 4837 | 4725 | JS | 4877 | 4572 | INCLUDE |  |  |  | 4837 | 4725 | 4781 | 4877 | 4572 | 4724.5 |  |  | 2 | 20 | 22
| 2229905006 |  |  | 2059 | 2185 | JS | 2146 | 2121 | INCLUDE |  |  |  | 2059 | 2185 | 2122 | 2146 | 2121 | 2133.5 |  |  | 1 | 10 | 11
| 2230003011 |  |  | 17052 | 18666 | JS | 16905 | 16248 | INCLUDE |  |  |  | 17052 | 18666 | 17859 | 16905 | 16248 | 16576.5 |  |  | 2 | 20 | 22 |  |  | x\y | 10 | 20 | 30 | 40 | 50 | Total | 1-k | 0
| 2230003937 |  |  | 7365 | 6763 | JS | 7259 | 6954 | INCLUDE |  |  |  | 7365 | 6763 | 7064 | 7259 | 6954 | 7106.5 |  |  | 2 | 20 | 22 |  |  | 1 | 0.46154 | 0 | 0 | 0 | 0 | 0.46154 | a | 1
| 2230106239 |  |  | 940 | 987 | JS | 996 | 977 | INCLUDE |  |  |  | 940 | 987 | 963.5 | 996 | 977 | 986.5 |  |  | 1 | 10 | 11 |  |  | 2 | 0 | 0.53846 | 0 | 0 | 0 | 0.53846 | b | 0
| 2230106308 |  |  | 434 | 383 | JS | 456 | 506 | INCLUDE |  |  |  | 434 | 383 | 408.5 | 456 | 506 | 481 |  |  | 1 | 10 | 11 |  |  | 3 | 0 | 0 | 0 | 0 | 0 | 0 | c | 1
| 2230104994 |  |  | 5693 | 5553 | JS | 5677 | 5529 | INCLUDE |  |  |  | 5693 | 5553 | 5623 | 5677 | 5529 | 5603 |  |  | 2 | 20 | 22 |  |  | 4 | 0 | 0 | 0 | 0 | 0 | 0 | se | 0
| 2230104292 |  |  | 4987 | 4860 | JS | 5071 | 5021 | INCLUDE |  |  |  | 4987 | 4860 | 4923.5 | 5071 | 5021 | 5046 |  |  | 2 | 20 | 22 |  |  | 5 | 0 | 0 | 0 | 0 | 0 | 0
| 2230106691 |  |  | 2754 | 2736 | JS | 2873 | 2726 | EXCLUDE | [16 empty] | Total | 0.46154 | 0.53846 | 0 | 0 | 0 | 1
| 2230204176 |  |  | 1725 | 1752 | JS | 1775 | 1668 | INCLUDE |  |  |  | 1725 | 1752 | 1738.5 | 1775 | 1668 | 1721.5 |  |  | 1 | 10 | 11
| 2231404848 |  |  | 7192 | 6922 | JS | 7454 | 7641 | INCLUDE |  |  |  | 7192 | 6922 | 7057 | 7454 | 7641 | 7547.5 |  |  | 2 | 20 | 22
| 2231406897 |  |  | 3024 | 2979 | JS | 3058 | 2961 | INCLUDE |  |  |  | 3024 | 2979 | 3001.5 | 3058 | 2961 | 3009.5 |  |  | 1 | 10 | 11 |  |  | Calc a | 0.46154
| 2231407330 |  |  | 903 | 888 | JS | 954 | 1019 | INCLUDE |  |  |  | 903 | 888 | 895.5 | 954 | 1019 | 986.5 |  |  | 1 | 10 | 11 |  |  |  |  | 0.53846
| 2231503109 |  |  | 11962 | 12349 | JS | 11793 | 11631 | INCLUDE |  |  |  | 11962 | 12349 | 12155.5 | 11793 | 11631 | 11712 |  |  | 2 | 20 | 22 | [5 empty] | 0
| 2231504117 |  |  | 1295 | 1286 | JS | 1347 | 1485 | INCLUDE |  |  |  | 1295 | 1286 | 1290.5 | 1347 | 1485 | 1416 |  |  | 1 | 10 | 11 | [6 empty] | 0
| 2231502835 |  |  | 948 | 938 | JS | 1000 | 997 | INCLUDE |  |  |  | 948 | 938 | 943 | 1000 | 997 | 998.5 |  |  | 1 | 10 | 11 | [7 empty] | 0
| 2231503694 |  |  | 11643 | 10982 | JS | 11982 | 11415 | INCLUDE |  |  |  | 11643 | 10982 | 11312.5 | 11982 | 11415 | 11698.5 |  |  | 2 | 20 | 22 |  |  | Calc b |  | 0 | 0 | 0 | 0
| 2231504139 |  |  | 2687 | 2713 | JS | 2777 | 2881 | INCLUDE |  |  |  | 2687 | 2713 | 2700 | 2777 | 2881 | 2829 |  |  | 1 | 10 | 11 |  |  |  | 0 |  | 0 | 0 | 0
| 2231503832 |  |  | 3688 | 3434 | JS | 3714 | 3839 | INCLUDE |  |  |  | 3688 | 3434 | 3561 | 3714 | 3839 | 3776.5 |  |  | 2 | 20 | 22 |  |  |  | 0 | 0 |  | 0 | 0
| 2231503992 |  |  | 1756 | 1666 | JS | 1873 | 1925 | INCLUDE |  |  |  | 1756 | 1666 | 1711 | 1873 | 1925 | 1899 |  |  | 1 | 10 | 11 |  |  |  | 0 | 0 | 0 |  | 0
| 2231506001 |  |  | 989 | 1023 | JS | 1062 | 1022 | INCLUDE |  |  |  | 989 | 1023 | 1006 | 1062 | 1022 | 1042 |  |  | 1 | 10 | 11 |  |  |  | 0 | 0 | 0 | 0
| 2231504468 |  |  | 1182 | 1179 | JS | 1570 | 1662 | INCLUDE |  |  |  | 1182 | 1179 | 1180.5 | 1570 | 1662 | 1616 |  |  | 1 | 10 | 11
| 2231602794 |  |  | 9098 | 8916 | JS | 9212 | 9549 | INCLUDE |  |  |  | 9098 | 8916 | 9007 | 9212 | 9549 | 9380.5 |  |  | 2 | 20 | 22
| 2231603495 |  |  | 4333 | 4543 | JS | 4427 | 4238 | INCLUDE |  |  |  | 4333 | 4543 | 4438 | 4427 | 4238 | 4332.5 |  |  | 2 | 20 | 22
| 2231503532 |  |  | 3486 | 3726 | JS | 3516 | 3101 | INCLUDE |  |  |  | 3486 | 3726 | 3606 | 3516 | 3101 | 3308.5 |  |  | 2 | 20 | 22
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
|  |  | x | y |  |  |  | Pearson Correlation Coefficient (R) | 0.99714
|  | N | 39 | 39 |  |  |  | Coefficient of Determination (R^2) | 0.99429
|  | Mean | 4572.1282 | 4614.4872
|  |  | Coefficents | LCL | UCL
Slope
Intercept
Concordance
n x n concordance table
|  | x\y | <3020 | ≥3020 |  |  |  | Total
|  | <3030 | 18 | 0 |  |  |  | 18
|  | ≥3030 | 0 | 21 |  |  |  | 21
|  | Total | 18 | 21 |  |  |  | 39
|  | Cohen's Kappa | 1 | Good agreement
|  | se | 0
|  | alpha | 0.05
|  | z-crit | 1.96
|  | CI lower | 1
|  | CI upper | 1
TRUENESS ASSESSED USING STANDARD MATERIALS
Trueness verification based on RCPAQAP material
EQA Program Name
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
| ENTER below MIMIMUM of 6 and MAXIMUM of 8 levels in single or replicate values for test method | [15 empty] | Calculations
[17 empty] | Trueness using RCPAQAP materials |  |  |  | Cut-off Values in reportable units
| Sample ID | Global Target | Method Group Median | First Replicate | Second Replicate | Mean | Trueness
[17 empty] | sd (sx) | s2x | Bias | %Bias | High cut off | Low cut off |  | n
| 23-08 | 200 | 195 | 190 | 190 | 190 | Pass | [9 empty] | 0 | 0 | -5 | -2.5641 | 186.7269 | 203.2731 |  | 2
| 23-09 | 250 | 230 | 232 | 232 | 232 | Pass | [9 empty] | 0 | 0 | 2 | 0.86957 | 220.2419 | 239.7581 |  | 2
| 23-10 | 10 | 20 | 15 |  | 15 | Fail | [9 empty] | 0 | 0 | -5 | -25 | 18.8 | 21.2 |  | 1
| 23-11 | 35 | 38 | 35 | 35 | 35 | Fail | [9 empty] | 0 | 0 | -3 | -7.8947 | 36.3878 | 39.6122 |  | 2
| 23-12 | 90 | 91 | 89 | 89 | 89 | Pass | [9 empty] | 0 | 0 | -2 | -2.1978 | 87.1392 | 94.8608 |  | 2
| 23-13 | 75 | 70 | 68 | 67 | 67.5 | Pass | [9 empty] | 0.70711 | 0.26726 | -2.5 | -3.5714 | 67.0302 | 72.9698 |  | 2
| 23-14 | 100 | 100 | 95 | 95 | 95 | Fail | [9 empty] | 0 | 0 | -5 | -5 | 95.7574 | 104.2426 |  | 2
| 23-15 | 200 | 205 | [13 empty] | 0 | 0 | [5 empty] | 0
Method Group Median CV from EQA Program
3%
| Number of labs in EQA Method Group |  |  |  |  | Std Err of method (sa) | 0.013416
| 5 |  |  |  |  | Number of samples tested | 7
[6 empty] | Number of replicates | 2
Trueness verification based on other reference materials
Trueness is verified if the TEST method gives results that fall within the calculated lower and upper verification limits.
This form of verification should be used if (1) the two methods are substantially different e.g. the test method uses a different
analytical principle to the reference method, or (2) the manufacturer has not provided relevant bias information, or (3) an
EQA program does not exist that can be used for trueness verification, or (4) alternative reference materials suitable for
trueness verification are available e.g. calibrator materials from different batch numbers. The default minimum acceptable
imprecision is as published by RCPAQAP.
Name of Material Tested
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
ENTER below MIMIMUM of 3 different batches of calibrators
| ID | Target Value | First Replicate | Second Replicate | Third Replicate | Mean | Strict | t-test
[17 empty] | SE | Abs Bias | Bias CI Lower | Bias CI Upper | Max Bias | Low cut off | High cut off
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
| Allowable bias |  | Alpha
5%
PRECISION
Material Used
Material Lot No.
Reagent Lot No.
Calibrator Lot No.
Claimed Imprecision
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Mean | 106 | 463 | 106 | 463
|  | Repeatability CV | 1% | 1% | 1% | 1%
|  | Within lab CV | 2.6% | 2.6% | 2.6% | 2.6%
|  | ENTER up to 5 values per day for each level | [18 empty] | Calculations
| Day | Run | Level 1 | Level 2 | Level 3 | Level 4 | [15 empty] | Level 1 | Level 2 | Level 3 | Level 4
| Day 1 | Run 1 | 95.1 | 418.2 | 95.1 | 418.2 | [13 empty] | sum | Day 1 | 477.4 | 2083.4 | 477.4 | 2083.4
| Day 1 | Run 1 | 95.2 | 415.2 | 95.2 | 415.2 | [14 empty] | Day 2 | 479.5 | 2086.1 | 479.5 | 2086.1
| Day 1 | Run 1 | 95.8 | 416.9 | 95.8 | 416.9 | [14 empty] | Day 3 | 465.5 | 2065.7 | 465.5 | 2065.7
| Day 1 | Run 1 | 95.4 | 416.3 | 95.4 | 416.3 | [14 empty] | Day 4 | 459.7 | 1999.6 | 459.7 | 1999.6
| Day 1 | Run 1 | 95.9 | 416.8 | 95.9 | 416.8 | [14 empty] | Day 5 | 456.9 | 2005.9 | 456.9 | 2005.9
| Day 2 | Run 1 | 98.2 | 421.3 | 98.2 | 421.3 | [13 empty] | n | Day 1 | 5 | 5 | 5 | 5
| Day 2 | Run 1 | 96.6 | 416.7 | 96.6 | 416.7 | [14 empty] | Day 2 | 5 | 5 | 5 | 5
| Day 2 | Run 1 | 96.9 | 422.5 | 96.9 | 422.5 | [14 empty] | Day 3 | 5 | 5 | 5 | 5
| Day 2 | Run 1 | 96.2 | 419.9 | 96.2 | 419.9 | [14 empty] | Day 4 | 5 | 5 | 5 | 5
| Day 2 | Run 1 | 91.6 | 405.7 | 91.6 | 405.7 | [14 empty] | Day 5 | 5 | 5 | 5 | 5
| Day 3 | Run 1 | 94 | 413.3 | 94 | 413.3 | [13 empty] | mean | Day 1 | 95.48 | 416.68 | 95.48 | 416.68
| Day 3 | Run 1 | 92 | 411.7 | 92 | 411.7 | [14 empty] | Day 2 | 95.9 | 417.22 | 95.9 | 417.22
| Day 3 | Run 1 | 93.1 | 413.8 | 93.1 | 413.8 | [14 empty] | Day 3 | 93.1 | 413.14 | 93.1 | 413.14
| Day 3 | Run 1 | 92.7 | 412.3 | 92.7 | 412.3 | [14 empty] | Day 4 | 91.94 | 399.92 | 91.94 | 399.92
| Day 3 | Run 1 | 93.7 | 414.6 | 93.7 | 414.6 | [14 empty] | Day 5 | 91.38 | 401.18 | 91.38 | 401.18
| Day 4 | Run 1 | 91.4 | 401.3 | 91.4 | 401.3 | [14 empty] | k | 5 | 5 | 5 | 5
| Day 4 | Run 1 | 91.6 | 398.7 | 91.6 | 398.7 | [14 empty] | N | 25 | 25 | 25 | 25
| Day 4 | Run 1 | 92.5 | 399 | 92.5 | 399 | [14 empty] | avg.n | 5 | 5 | 5 | 5
| Day 4 | Run 1 | 91.2 | 398.2 | 91.2 | 398.2 | [14 empty] | mean | 93.56 | 409.628 | 93.56 | 409.628
| Day 4 | Run 1 | 93 | 402.4 | 93 | 402.4 | [14 empty] | SS.total | 144.58 | 1653.3904 | 144.58 | 1653.3904
| Day 5 | Run 1 | 92.9 | 401.4 | 92.9 | 401.4 | [14 empty] | SS1 | 83.752 | 1426.5864 | 83.752 | 1426.5864
| Day 5 | Run 1 | 92.1 | 402 | 92.1 | 402 | [14 empty] | SS2 | 60.828 | 226.804 | 60.828 | 226.804
| Day 5 | Run 1 | 94.7 | 404.2 | 94.7 | 404.2 | [14 empty] | DF.total | 24 | 24 | 24 | 24
| Day 5 | Run 1 | 89.2 | 399.7 | 89.2 | 399.7 | [14 empty] | DF1 | 4 | 4 | 4 | 4
| Day 5 | Run 1 | 88 | 398.6 | 88 | 398.6 | [14 empty] | DF2 | 20 | 20 | 20 | 20
[21 empty] | MS1 | 20.938 | 356.6466 | 20.938 | 356.6466
[21 empty] | MS2 | 3.0414 | 11.3402 | 3.0414 | 11.3402
|  | Repeatability | 1.744 | 3.3675 | 1.744 | 3.3675 | [14 empty] | n^2 | 25 | 25 | 25 | 25
|  | Between day | 1.8919 | 8.3103 | 1.8919 | 8.3103 | [15 empty] | 25 | 25 | 25 | 25
|  | Within lab | 2.5731 | 8.9667 | 2.5731 | 8.9667 | [15 empty] | 25 | 25 | 25 | 25
[22 empty] | 25 | 25 | 25 | 25
|  | Repeatability CV | 1.864% | 0.8221% | 1.864% | 0.8221% | [15 empty] | 25 | 25 | 25 | 25
|  | Between day CV | 2.022% | 2.029% | 2.022% | 2.029% | [14 empty] | SN2 | 125 | 125 | 125 | 125
|  | Within lab CV | 2.75% | 2.189% | 2.75% | 2.189% | [14 empty] | n.0 | 5 | 5 | 5 | 5
[21 empty] | a.1 | 0.2 | 0.2 | 0.2 | 0.2
[21 empty] | a.2 | 0.8 | 0.8 | 0.8 | 0.8
|  |  | Upper Verification Limits | [17 empty] | den1 | 4.384 | 1271.968 | 4.384 | 1271.968
|  | Repeatability UVL | 1.356% | 1.356% | 1.356% | 1.356% | [14 empty] | den2 | 0.296 | 4.1152 | 0.296 | 4.1152
|  | Within lab UVL | 3.896% | 4.405% | 3.896% | 4.405% | [14 empty] | num | 43.8339 | 6464.398 | 43.8339 | 6464.398
[21 empty] | DFwl | 9.3662 | 5.0658 | 9.3662 | 5.0658
|  |  | Comparison to Claim | [17 empty] | nLevels | 4
|  |  | Level 1 | Level 2 | Level 3 | Level 4 | [14 empty] | alpha | 0.05
|  | Repeatability | Fail | Pass | Fail | Pass | [14 empty] | s_r | 1.744 | 3.3675 | 1.744 | 3.3675
|  | Within Lab Precision | Pass | Pass | Pass | Pass | [14 empty] | s_bd | 1.8919 | 8.3103 | 1.8919 | 8.3103
[21 empty] | s_wl | 2.5731 | 8.9667 | 2.5731 | 8.9667
[21 empty] | Claimed s_r | 0.9356 | 4.0963 | 0.9356 | 4.0963
|  |  | Assay Capability | [17 empty] | Claimed s_wl | 2.4326 | 10.6503 | 2.4326 | 10.6503
|  |  | Level 1 | Level 2 | Level 3 | Level 4 | [13 empty] | ꭕ2 within group (Analyse-IT protocol) | ꭕ2 _r | 69.4901 | 13.5167 | 69.4901 | 13.5167
|  | Mean | 93.56 | 409.628 | 93.56 | 409.628 | [13 empty] | P value within group (Analyse-IT protocol) | p_r | 2.2057e-07 | 0.85413 | 2.2057e-07 | 0.85413
|  | RCPA APS | 10.69% | 10% | 10.69% | 10% | [13 empty] | ꭕ2 within laboratory (Analyse-IT Protocol) | ꭕ2 _wl | 10.4795 | 3.5908 | 10.4795 | 3.5908
|  | Capability (σ) | 3.8864 | 4.5683 | 3.8864 | 4.5683 | [13 empty] | P value within lab (Analyse-IT protocol) | p_wl | 0.31307 | 0.6097 | 0.31307 | 0.6097
|  | Acceptability | Poor | Acceptable | Poor | Acceptable | [13 empty] | ꭕ2 within group (CLSI EP15 protocol) | ꭕ2 _r | 36.7605 | 36.7605 | 36.7605 | 36.7605
[20 empty] | ꭕ2 within laboratory (CLSI EP15 Protocol) | ꭕ2 wl | 21.0341 | 14.5439 | 21.0341 | 14.5439
[20 empty] | F statistic within group | F_r | 1.3557 | 1.3557 | 1.3557 | 1.3557
|  |  | Comparison to Biological Variation APS | [16 empty] | F statistic within laboratory | F_WL | 1.4986 | 1.6944 | 1.4986 | 1.6944
|  | CVi | 4% |  |  | Specification | [14 empty] | rho | 2.6 | 2.6 | 2.6 | 2.6
|  | CVg | 7.9% |  | Imprecision | 2%
|  | Required Category | Desirable |  | Bias | 2.214%
|  | alpha | 0.05 |  | Total Error | 5.503%
Upper Verification Limits
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Repeatability | 2.711% | 2.711% | 2.711% | 2.711%
|  | Within Lab Precision | 2.997% | 3.389% | 2.997% | 3.389%
Imprecision Verification Result
|  | Repeatability | Pass | Pass | Pass | Pass
|  | Within Lab Precision | Pass | Pass | Pass | Pass
REFERENCE INTERVALS
Reference Intervals that will be used
Reference Interval Verification
| LRL | URL |  |  |  | LRL | URL
| 30 | 110 |  |  |  | 30 | 110
| Sample ID | Operator | Result |  |  | Sample ID | Operator | Result
1 | Test 1 | T1 | 61 | 0 |  | Test 1 | T1 | 61 | 0
2 | Test 2 | T2 | 97 | 0 |  | Test 2 | T2 | 97 | 0
3 | Test 3 | T3 | 93 | 0 |  | Test 3 | T3 | 93 | 0
4 | Test 4 | T4 | 48 | 0 |  | Test 4 | T4 | 48 | 0
5 | Test 5 | T5 | 52 | 0 |  | Test 5 | T5 | 52 | 0
6 | Test 6 | T6 | 52 | 0 |  | Test 6 | T6 | 52 | 0
7 | Test 7 | T7 | 41 | 0 |  | Test 7 | T7 | 41 | 0
8 | Test 8 | T8 | 75 | 0 |  | Test 8 | T8 | 75 | 0
9 | Test 9 | T9 | 123 | 1 |  | Test 9 | T9 | 123 | 1
10 | Test 10 | T10 | 55 | 0 |  | Test 10 | T10 | 55 | 0
11 | Test 11 | T11 | 56 | 0 |  | Test 11 | T11 | 56 | 0
12 | Test 12 | T12 | 111 | 1 |  | Test 12 | T12 | 111 | 1
13 | Test 13 | T13 | 74 | 0 |  | Test 13 | T13 | 74 | 0
14 | Test 14 | T14 | 73 | 0 |  | Test 14 | T14 | 73 | 0
15 | Test 15 | T15 | 67 | 0 |  | Test 15 | T15 | 67 | 0
16 | Test 16 | T16 | 95 | 0 |  | Test 16 | T16 | 95 | 0
17 | Test 17 | T17 | 55 | 0 |  | Test 17 | T17 | 55 | 0
18 | Test 18 | T18 | 44 | 0 |  | Test 18 | T18 | 44 | 0
19 | Test 19 | T19 | 59 | 0 |  | Test 19 | T19 | 59 | 0
20 | Test 20 | T20 | 112 | 1 |  | Test 20 | T20 | 112 | 1
21 | [5 empty] | Test 21 | T21 | 95 | 0
22 | [5 empty] | Test 22 | T22 | 90 | 0
23 | [5 empty] | Test 23 | T23 | 44 | 0
24 | [5 empty] | Test 24 | T24 | 65 | 0
25 | [5 empty] | Test 25 | T25 | 62 | 0
26 | [5 empty] | Test 26 | T26 | 46 | 0
27 | [5 empty] | Test 27 | T27 | 97 | 0
28 | [5 empty] | Test 28 | T28 | 56 | 0
29 | [5 empty] | Test 29 | T29 | 70 | 0
30 | [5 empty] | Test 30 | T30 | 71 | 0
31 | [5 empty] | Test 31 | T31 | 48 | 0
32 | [5 empty] | Test 32 | T32 | 64 | 0
33 | [5 empty] | Test 33 | T33 | 44 | 0
34 | [5 empty] | Test 34 | T34 | 75 | 0
35 | [5 empty] | Test 35 | T35 | 57 | 0
36 | [5 empty] | Test 36 | T36 | 52 | 0
37 | [5 empty] | Test 37 | T37 | 69 | 0
38 | [5 empty] | Test 38 | T38 | 95 | 0
39 | [5 empty] | Test 39 | T39 | 51 | 0
40 | [5 empty] | Test 40 | T40 | 74 | 0
|  | Reference interval verification | Fail |  |  |  | Reference interval verification | Pass
|  | N | 20 |  |  |  | N | 40
|  | Outliers | 3 |  |  |  | Outliers | 3
|  | p | 0.015902 |  |  |  | p | 0.13815
LIMIT OF BLANK, DETECTION, QUANTITATION
| Date of analysis |  |  | Reagent Lot No.
| Operator ID |  |  | Calibrator Lot No.
Limit of Blank (LoB)
| ENTER the LoB claim |  | 0.9 | ng/L
ENTER min of 2 replicates, of min 2 blank samples, of min 3 days, for min of 20 values
[9 empty] | Bounds for observed proportion of results relative to a detection capability claim | [10 empty] | Values ≤ LoB claim
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Total number of measurements (n) |  |  | % Observed proportion boundary | [7 empty] | 1 | 1 | 1
|  | Day 1 | 2022-10-23 | 0.5 | 0.8 |  |  |  | 20 |  |  | 85% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.5 | 0.5 |  |  |  | 30 |  |  | 86.67% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.6 | 0.6 |  |  |  | 40 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 2 | 2022-10-24 | 0.7 | 0.7 |  |  |  | 50 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.7 | 0.7 |  |  |  | 60 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.8 | 0.9 |  |  |  | 70 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 3 | 2022-10-25 | 0.5 | 0.6 |  |  |  | 80 |  |  | 91.25% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.4 | 0.5 |  |  |  | 90 |  |  | 91.11% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.6 | 0.7 |  |  |  | 100 |  |  | 91% |  |  |  |  | n <= LoB claim |  | 27
[9 empty] | 150 |  |  | 92% |  |  |  |  | % n<= LoB claim |  | 150 | %
[9 empty] | 200 |  |  | 92.5%
| Total number of values |  | 18 |  | Min Req | 15 |  |  | 250 |  |  | 92.8%
| mean of blank |  | 0.62778 |  | % Min Req | 83.33% |  |  | 300 |  |  | 93%
| SD of blank |  | 0.12826 | [5 empty] | 400 |  |  | 93.25%
Limit of blank (LoB) |  |  | 0.84232 | 0 | Alpha | 0.05 |  |  | 500 |  |  | 93.4%
|  |  | Pass |  | C_beta | 1.6727 |  |  | 1000 |  |  | 93.8%
Limit of Detection (LoD)
| ENTER the LoD claim |  | 1.1 | ng/L
Values ≥ LoB claim
|  | ENTER min of 2 replicates, of min 2 LoD claim samples, of min 3 days, for min of 20 values | [17 empty] | 1 | 1
1
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  | Selected samples should have a concentration at the claimed LoD | [11 empty] | 1 | 1
|  | Day 1 | 2022-10-23 | 1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 0.9 | 0.8 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 1 | 1.1 | [14 empty] | 1 | 1
|  | Day 2 | 2022-10-24 | 1.1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 0.9 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 1.1 | 1.1 | [14 empty] | 1 | 1
|  | Day 3 | 2022-10-25 | 1.2 | 1.1 | [11 empty] | n >=LoB claim |  | 17
|  |  | 2022-10-25 | 1.1 | 0.9 | [11 empty] | % n >= LoB claim |  | 94.4444 | %
|  |  | 2022-10-25 | 1 | 1.1
|  |  | Sample SD | 0.094281 | 0.099381
|  |  |  | 0.0088889 | 0.0098765 | #N/A
| Total number of values |  | 18 |  | Min Req | 15
| Pooled SD |  | 0.096864 |  | % Min Req | 83.33%
|  | Limit of Detection | 1.062 | 0
|  |  | Pass |  | Alpha | 0.05
[5 empty] | C_beta | 1.6727
Limit of Quantitaion (LoQ)
| Enter the LoQ claim |  | 1.5 | ng/L
ENTER min of 2 replicates, of min 2 LoQ claim samples, of min 3 days, for min of 20 values
% difference from means
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Sample 1 | Sample 2 | Sample 3 |  |  |  |  | Values < TEa |  |  |  | % difference calculations
|  | Day 1 | 2022-10-23 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-23 | 13 | 14 |  |  |  | 14.5985 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 14.5985 | 8.0292 | #VALUE!
|  |  | 2022-10-23 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  | Day 2 | 2022-10-24 | 16 | 15 |  |  |  | 5.1095 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 5.1095 | 1.4599 | #VALUE!
|  |  | 2022-10-24 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-24 | 17 | 14 |  |  |  | 11.6788 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 11.6788 | 8.0292 | #VALUE!
|  | Day 3 | 2022-10-25 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-25 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-25 | 18 | 16 |  |  |  | 18.2482 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 18.2482 | 5.1095 | #VALUE!
|  | Sample mean |  | 15.2222 | 15.2222 | [7 empty] | n >= TEa |  | 0
|  |  | Sample SD | 1.4741 | 0.78567 | [7 empty] | % n >= TEa |  | 0
|  |  | Sample CV | 9.684% | 5.161%
| Total number of values |  | 18 |  | Min Req | 15
| Total number of samples |  | 2 |  | % Min Req | 83.33%
Grand Mean |  |  | 15.2222
[5 empty] | Alpha | 0.05
| TEa |  | #REF! | % | C_beta | 1.6727
Fail
Note on B380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on C380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit
Note on G380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on H380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit

## APS
Description | Sample Type | Units | CVi | CVg | APS Relative | APS Absolute | APS Cut | APS Src
17-Hydroxyprogesterone | Serum | nmol/L | 19.6% | 52.4% | 20% | 2 | 10 | RCPA QAP
Acid Phosphatase (Prostatic) | Serum | U/L | 33.8% |  | 15% | 1.5 | 10 | RCPA QAP
Acid Phosphatase (Tartrate Resistant) | Serum | U/L | 8% | 13.3% | 15% | 1.5 | 10 | RCPA QAP
Acid Phosphatase (Total) | Serum | U/L | 8.9% | 8% | 15% | 1.5 | 10 | RCPA QAP
Activated Partial Thromboplastin Time | Plasma | s | 2.7% | 8.6% | 25% | 10 | 40 | RCPA QAP
Adrenaline | Plasma | pmol/L | 48.3% |  | 30% | 30 | 100 | RCPA QAP
Adrenaline (Urine) | Urine | nmol/L |  |  | 30% | 30 | 100 | RCPA QAP
Adrenocorticotropic hormone (mass) | Plasma | ng/L |  |  | 10% | 9 | 90 | RCPA QAP
Adrenocorticotropic hormone (molar) | Plasma | pmol/L |  |  | 10% | 2 | 20 | RCPA QAP
Alanine aminotransferase | Serum/Plasma | U/L | 9.6% | 28% | 12% | 5 | 40 | RCPA QAP
Albumin | Serum/Plasma | g/L | 2.5% | 5.1% | 6% | 2 | 33 | RCPA QAP
Albumin (CSF) | CSF | g/L |  |  | 20% | 0.02 | 0.1 | RCPA QAP
Albumin (Urine) | Urine | mg/L | 36% | 55% | 20% | 4 | 20 | RCPA QAP
Aldosterone | Serum/Plasma | pmol/L | 29.4% | 40.1% | 15% | 24 | 160 | RCPA QAP
Alkaline Phosphatase | Serum/Plasma | U/L | 6.4% | 24.8% | 12% | 15 | 125 | RCPA QAP
Alphafetoprotein | Serum/Plasma | kIU/L | 12% | 46% | 12% | 2 | 17 | RCPA QAP
Amikacin | Serum/Plasma | mg/L |  |  | 10% | 2 | 20 | RCPA QAP
Aminolevulinic acid | Urine | umol/L |  |  | 25% | 8 | 32 | RCPA QAP
Amiodarone | Serum/Plasma | mg/L |  |  | 10% | 0.13 | 1.3 | RCPA QAP
Amitriptyline | Serum/Plasma | ug/L |  |  | 10% | 3 | 28 | RCPA QAP
Ammonia | Serum/Plasma | umol/L |  |  | 10% | 5 | 50 | RCPA QAP
Amphetamine (class) (Urine) | Urine | ug/L |  |  | 25% | 75 | 300 | RCPA QAP
Amylase | Serum/Plasma | U/L | 6.6% | 30.1% | 10% | 10 | 100 | RCPA QAP
Amylase (Urine) | Urine | U/L | 94% | 46% | 20% | 20 | 100 | RCPA QAP
Androstenedione | Serum/Plasma | nmol/L | 11.1% | 51.1% | 15% | 1.5 | 10 | RCPA QAP
Angiotensin Converting Enzyme | Plasma | U/L | 12.5% | 27.7% | 10% | 1 | 10 | RCPA QAP
Ant Mullerian Hormone | Serum/Plasma | pmol/L | 19.2% | 20.4% | 15% | 1 | 6 | RCPA QAP
Apolipoprotein A1 | Serum/Plasma | g/L | 5.8% | 11.2% | 10% | 0.2 | 2 | RCPA QAP
Apolipoprotein B | Serum/Plasma | g/L | 7.4% | 20.2% | 10% | 0.2 | 2 | RCPA QAP
Aspartate aminotransferase | Serum/Plasma | U/L | 9.5% | 21% | 12% | 5 | 40 | RCPA QAP
B-2-Microglobulin | Serum/Plasma | mg/L | 4.1% | 11.5% | 10% | 0.2 | 2 | RCPA QAP
Barbiturates (Urine) | Urine | mg/L |  |  | 25% | 50 | 200 | RCPA QAP
Benzodiazepines (class) (Urine) | Urine | ug/L |  |  | 25% | 50 | 200 | RCPA QAP
Beta Hydroxybutyrate (mmol/L) | Serum/Plasma | mmol/L | 4% | 19.5% | 15% | 0 | 0 | RCPA QAP
Beta Hydroxybutyrate (umol/L) | Serum/Plasma | umol/L | 4% | 19.5% | 15% | 0 | 0 | RCPA QAP
Bicarbonate | Serum/Plasma | mmol/L | 4.8% | 4.7% | 10% | 2 | 20 | RCPA QAP
Bile Acids (Total) | Serum/Plasma | umol/L |  |  | 10% | 4 | 40 | RCPA QAP
Bilirubin (Conjugated) | Serum/Plasma | umol/L | 36.8% | 43.2% | 20% | 3 | 15 | RCPA QAP
Bilirubin (CSF) | CSF | umol/L |  |  | 20% | 0.12 | 0.6 | RCPA QAP
Bilirubin (Neonatal) | Serum/Plasma | umol/L | 23.8% | 39% | 10% | 8 | 80 | RCPA QAP
Bilirubin (Total) | Serum/Plasma | umol/L | 23.8% | 39% | 12% | 3 | 25 | RCPA QAP
Blood (Faeces) | Faeces | ug/L |  |  | 15% | 15 | 100 | RCPA QAP
Brain Natriuretic Peptide | Plasma | ng/L | 44% | 56% | 20% | 20 | 100 | RCPA QAP
C3 Complement | Serum/Plasma | g/L | 5.2% | 15.6% | 10% | 0.1 | 1 | RCPA QAP
C4 Complement | Serum/Plasma | g/L | 8.9% | 33.4% | 10% | 0.02 | 0.2 | RCPA QAP
CA 125 | Serum/Plasma | kU/L | 24.7% | 54.6% | 12% | 6 | 50 | RCPA QAP
CA 15-3 | Serum/Plasma | kU/L | 6.1% | 62.9% | 12% | 6 | 50 | RCPA QAP
CA 19-9 | Serum/Plasma | kU/L | 22.5% | 82.4% | 15% | 6 | 40 | RCPA QAP
Caeruloplasmin | Serum/Plasma | g/L | 5.8% | 11.1% | 15% | 0.03 | 0.2 | RCPA QAP
Calcitonin | Serum/Plasma | ng/L |  |  | 10% | 2 | 20 | RCPA QAP
Calcium | Serum/Plasma | mmol/L | 1.9% | 2.8% | 4% | 0.1 | 2.5 | RCPA QAP
Calcium (Ionised) | Blood | mmol/L | 1.7% | 1.9% | 4% | 0.04 | 1 | RCPA QAP
Calcium (Urine) | Urine | mmol/L | 27.5% | 36.6% | 10% | 0.2 | 2 | RCPA QAP
Cannabinoids (class) (Urine) | Urine | ug/L |  |  | 25% | 12.5 | 50 | RCPA QAP
Carbamazepine (mass) | Serum/Plasma | mg/L |  |  | 10% | 0.5 | 4.7 | RCPA QAP
Carboxyhaemoglobin | Blood | % |  |  | 20% | 1 | 5 | RCPA QAP
Carcinoembryonic Antigen | Serum/Plasma | ug/L | 12.7% | 55.6% | 12% | 0.6 | 5 | RCPA QAP
Chloride | Serum/Plasma | mmol/L | 1.1% | 1.3% | 3% | 3 | 100 | RCPA QAP
Chloride (Blood) | Blood | mmol/L | 1.1% | 1.3% | 3% | 3 | 100 | RCPA QAP
Chloride (Sweat) | Sweat | mmol/L | 15% | 25% | 5% | 2 | 40 | RCPA QAP
Chloride (Urine) | Urine | mmol/L | 24% | 16.7% | 10% | 4 | 40 | RCPA QAP
Cholesterol | Serum/Plasma | mmol/L | 5.5% | 15.8% | 6% | 0.3 | 5 | RCPA QAP
Cholinesterase | Serum/Plasma | kU/L | 7% | 10.4% | 10% | 500 | 5000 | RCPA QAP
Ciclosporin (Blood) | Blood | ug/L |  |  | 10% | 10 | 100 | RCPA QAP
CO2 (Partial pressure) | Blood | mmHg | 4.8% | 5.3% | 60% | 2 | 34 | RCPA QAP
Cocaine (class) (Urine) | Urine | ug/L |  |  | 25% | 75 | 300 | RCPA QAP
Conductivity | Sweat | mmol/L |  |  | 5% | 2 | 50 | RCPA QAP
Copper | Serum/Plasma | umol/L | 8% | 19% | 8% | 1.6 | 20 | RCPA QAP
Cortisol | Serum/Plasma | nmol/L | 22.7% | 31.9% | 15% | 15 | 100 | RCPA QAP
Cortisol (Saliva) | Saliva | nmol/L | 32.1% | 38.1% | 15% | 1 | 6 | RCPA QAP
Cortisol (Urine) | Urine | nmol/L |  |  | 15% | 30 | 200 | RCPA QAP
C-Peptide | Serum/Plasma | nmol/L | 16.6% | 23.2% | 12% | 0.15 | 1.25 | RCPA QAP
C-Reactive Protein | Serum/Plasma | mg/L | 42.2% | 76.3% | 20% | 0.8 | 4 | RCPA QAP
C-Reactive Protein (high sensitivity) | Serum/Plasma | mg/L | 42.2% | 76.3% | 20% | 0.6 | 3 | RCPA QAP
Creatine Kinase | Serum/Plasma | U/L | 15.4% | 31.9% | 12% | 15 | 125 | RCPA QAP
Creatine Kinase MB | Serum/Plasma | ug/L | 18.4% | 61.2% | 20% | 3 | 15 | RCPA QAP
Creatinine | Serum/Plasma | umol/L | 4.5% | 13.9% | 8% | 8 | 100 | RCPA QAP
Creatinine (Blood) | Blood | umol/L | 4.5% | 13.9% | 8% | 8 | 100 | RCPA QAP
Creatinine (enzymatic) | Serum/Plasma | umol/L | 4.5% | 13.9% | 8% | 8 | 100 | RCPA QAP
Creatinine (Urine) | Urine | mmol/L | 24% | 24.5% | 10% | 0.5 | 5 | RCPA QAP
CTX Beta-Crosslaps | Serum/Plasma | ng/L | 15.1% | 45.9% | 10% | 10 | 100 | RCPA QAP
D-Dimer | Plasma | mg/L | 23.3% | 26.5% | 50% | 0.5 | 1 | RCPA QAP
Dehydroepiandrosterone Sulfate | Serum/Plasma | umol/L | 7.1% | 21% | 12% | 1.2 | 10 | RCPA QAP
Deoxypyridinoline crosslinks (Urine) | Urine | nmol/L | 13.1% | 19% | 15% | 0 | 0 | State of Art
Digoxin (mass) | Serum/Plasma | ug/L |  |  | 10% | 0.2 | 1.6 | RCPA QAP
Dihydrotestosterone | Serum/Plasma | nmol/L |  |  | 15% | 0.1 | 0.5 | RCPA QAP
Dopamine (Urine) | Urine | umol/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Ethanol | Serum/Plasma | mmol/L |  |  | 10% | 2 | 20 | RCPA QAP
Everolimus | Serum/Plasma | ug/L |  |  | 10% | 0.5 | 5 | RCPA QAP
Ferritin | Serum/Plasma | ug/L | 14.9% | 13.7% | 15% | 2 | 27 | RCPA QAP
Fibrinogen | Plasma | g/L | 10.7% | 15.8% | 25% | 0.2 | 1 | RCPA QAP
Folate | Serum/Plasma | nmol/L | 12% | 66% | 25% | 1.5 | 6 | RCPA QAP
Folate (Red Cell) | Blood | nmol/L |  |  | 20% | 80 | 400 | RCPA QAP
Follicle Stimulating Hormone | Serum/Plasma | IU/L | 16% | 46.4% | 10% | 1 | 10 | RCPA QAP
Free Beta Human Chorionic Gonadotropin | Serum/Plasma | IU/L | 3.8% | 4.5% | 10% | 1 | 10 | Biological Variation
Fructosamine | Serum/Plasma | umol/L | 2.3% | 6.3% | 6% | 15 | 250 | RCPA QAP
Gamma glutamyltransferase | Serum/Plasma | U/L | 9.1% | 44.5% | 12% | 5 | 40 | RCPA QAP
Gastrin | Serum/Plasma | pmol/L |  |  | 10% | 10 | 100 | RCPA QAP
Gentamicin | Serum/Plasma | mg/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Glucose | Serum/Plasma | mmol/L | 5% | 8.1% | 8% | 0.4 | 5 | RCPA QAP
Glucose (Blood) | Blood | mmol/L | 5% | 8.1% | 8% | 0.4 | 5 | RCPA QAP
Glucose (CSF) | CSF | mmol/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Glucose (Urine) | Urine | mmol/L |  |  | 10% | 1 | 10 | RCPA QAP
Growth Hormone | Serum/Plasma | mIU/L |  |  | 15% | 1 | 7 | RCPA QAP
Haematocrit (POCT) | Blood | % | 2.3% | 5.5% | 20% | 4 | 20 | RCPA QAP
Haemoglobin (Plasma Free) | Serum/Plasma | g/L | 8.71% | 7.1% | 10% | 0.08 | 0.8 | RCPA QAP
Haemoglobin (POCT) | Blood | g/L | 3% | 7% | 5% | 5 | 100 | RCPA QAP
Haptoglobin | Serum/Plasma | g/L | 20.4% | 36.4% | 10% | 0.1 | 0.8 | RCPA QAP
HbA1c (IFCC) | Blood | mmol/mol | 1.6% | 7.1% | 8% | 4 | 45 | RCPA QAP
HbA1c (NGSP) | Blood | % | 1.6% | 7.1% | 6% | 0.4 | 6.7 | RCPA QAP
HDL Cholesterol | Serum/Plasma | mmol/L | 7.5% | 23% | 12% | 0.1 | 0.8 | RCPA QAP
Holotranscobalamin | Serum | pmol/L | 16% | 35% | 15% | 5 | 35 | RCPA QAP
Homocysteine | Serum/Plasma | umol/L | 9% | 40.3% | 10% | 1.5 | 15 | RCPA QAP
Homovanillic Acid (Urine) | Urine | umol/L | 41% | 71.3% | 15% | 6 | 40 | RCPA QAP
Human Chorionic Gonadotropin | Serum/Plasma | U/L |  |  | 10% | 1 | 10 | RCPA QAP
Human Epididymal Protein 4 | Serum/Plasma | pmol/L | 7.1% | 17.9% | 15% | 10 | 150 | RCPA QAP
Hydroxyindole Acetic Acid | Urine | umol/L |  |  | 20% | 8 | 40 | RCPA QAP
IGF Binding Protein 3 | Serum/Plasma | mg/L | 10.1% | 63.9% | 15% | 0.46 | 2.88 | RCPA QAP
Immunoglobulin A | Serum/Plasma | g/L | 5.7% | 19.6% | 10% | 0.06 | 0.6 | RCPA QAP
Immunoglobulin G | Serum/Plasma | g/L | 3.5% | 17.1% | 10% | 0.6 | 6 | RCPA QAP
Immunoglobulin G (CSF) | CSF | g/L |  |  | 20% | 0.02 | 0.1 | RCPA QAP
Immunoglobulin M | Serum/Plasma | g/L | 5.9% | 48.5% | 15% | 0.08 | 0.5 | RCPA QAP
Insulin | Serum/Plasma | mIU/L | 37.1% | 81.8% | 12% | 0.6 | 5 | RCPA QAP
Insulin Like Growth Factor 1 (mass) | Serum/Plasma | ug/L | 9.4% | 27% | 12% | 23 | 191 | RCPA QAP
Insulin Like Growth Factor 1 (molar) | Serum/Plasma | nmol/L | 9.4% | 27% | 12% | 3 | 25 | RCPA QAP
Iron | Serum/Plasma | umol/L | 26.5% | 23.2% | 12% | 3 | 23 | RCPA QAP
Itraconazole | Serum/Plasma | mg/L |  |  | 10% | 0.1 | 1 | RCPA QAP
Lactate | Serum/Plasma | mmol/L | 27.2% | 16% | 12% | 0.5 | 4 | RCPA QAP
Lactate (Blood) | Blood | mmol/L | 27.2% | 16.7% | 12% | 0.5 | 4 | RCPA QAP
Lactate (CSF) | CSF | mmol/L |  |  | 10% | 0.3 | 3 | RCPA QAP
Lactate Dehydrogenase | Serum/Plasma | U/L | 5.2% | 11.9% | 8% | 20 | 250 | RCPA QAP
LDL Cholesterol | Serum/Plasma | mmol/L | 8% | 27% | 10% | 0.2 | 2 | RCPA QAP
Lipase | Serum/Plasma | U/L | 7.7% | 23.4% | 20% | 12 | 60 | RCPA QAP
Lipoprotein (a) (mass) | Serum/Plasma | g/L | 7.6% | 135.8% | 20% | 0.06 | 0.3 | RCPA QAP
Lithium | Serum/Plasma | mmol/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Luteinising Hormon | Serum/Plasma | IU/L | 23.1% | 27.7% | 15% | 1.5 | 10 | RCPA QAP
Magnesium | Serum/Plasma | mmol/L | 3.6% | 6.4% | 8% | 0.1 | 1.25 | RCPA QAP
Magnesium (Urine) | Urine | mmol/L | 45.4% | 37.4% | 15% | 0.3 | 2 | RCPA QAP
Metanephrine (Plasma) | Plasma | pmol/L |  |  | 15% | 75 | 500 | RCPA QAP
Metanephrine (Urine) | Urine | umol/L | 46.3% | 76.8% | 20% | 0.2 | 1 | RCPA QAP
Methadone Metabolite | Urine | ug/L |  |  | 25% | 25 | 100 | RCPA QAP
Methaemoglobin (Blood) | Blood | % |  |  | 20% | 2 | 10 | RCPA QAP
Methotrexate | Serum/Plasma | umol/L |  |  | 10% | 0.1 | 1 | RCPA QAP
Methoxytyramine (Plasma) | Plasma | pmol/L |  |  | 30% | 36 | 120 | RCPA QAP
Methoxytyramine (Urine) | Urine | nmol/L |  |  | 15% | 15 | 100 | RCPA QAP
Mycophenolate | Serum/Plasma | mg/L |  |  | 10% | 0.1 | 1 | RCPA QAP
Neuron Specific Enolase | Serum/Plasma | ug/L | 10.9% | 16.6% | 10% | 2 | 20 | RCPA QAP
Normetanephrine (Plasma) | Plasma | pmol/L |  |  | 20% | 200 | 1000 | RCPA QAP
Normetanephrine (Urine) | Urine | umol/L |  |  | 20% | 0.4 | 2 | RCPA QAP
Nortriptyline | Serum/Plasma | ug/L |  |  | 10% | 3 | 26 | RCPA QAP
N-telopeptide (Urine) | Urine | nM BCE/L | 17% | 52% | 10% | 80 | 700 | RCPA QAP
N-terminal proBNP | Serum/Plasma | ng/L | 20.8% | 36.2% | 20% | 25 | 125 | RCPA QAP
Oestradiol | Serum/Plasma | pmol/L | 14.6% | 13.6% | 25% | 25 | 100 | RCPA QAP
Oestriol | Serum/Plasma | nmol/L |  |  | 15% | 0.9 | 6 | RCPA QAP
Opiates (class) (Urine) | Urine | ug/L |  |  | 25% | 75 | 300 | RCPA QAP
Osmolality | Serum/Plasma | mmol/kg | 0.8% | 0.9% | 3% | 8 | 266 | RCPA QAP
Osmolality (Urine) | Urine | mmol/kg | 28.3% | 57.9% | 2% | 6 | 300 | RCPA QAP
Oxygen partial pressure (Blood) | Blood | mmHg |  |  | 6% | 2 | 34 | RCPA QAP
Oxyhaemoglobin (Blood) | Blood | % |  |  | 4% | 3 | 75 | RCPA QAP
Paracetamol (mass) | Serum/Plasma | mg/L |  |  | 10% | 3 | 30 | RCPA QAP
Paracetamol (molar) | Serum/Plasma | umol/L |  |  | 10% | 20 | 200 | RCPA QAP
Parathyroid Hormone (mass) | Serum | ng/L | 15.7% | 23.5% | 12% | 9.43 | 75.4 | RCPA QAP
Parathyroid Hormone (molar) | Serum/Plasma | pmol/L | 15.7% | 23.5% | 12% | 1 | 8 | RCPA QAP
pH (Blood) | Blood | - | 0.2% | 0% | 0% | 0.04 | 50 | RCPA QAP
pH (Fluid) | Body Fluid | - |  |  | 0% | 0.2 | 50 | RCPA QAP
Phenobarbitone (mass) | Serum/Plasma | mg/L |  |  | 10% | 0.7 | 7 | RCPA QAP
Phenobarbitone (molar) | Serum/Plasma | umol/L |  |  | 10% | 3 | 30 | RCPA QAP
Phenytoin (mass) | Serum/Plasma | mg/L |  |  | 10% | 0.8 | 7.6 | RCPA QAP
Phenytoin (molar) | Serum/Plasma | umol/L |  |  | 10% | 3 | 30 | RCPA QAP
Phosphate | Serum/Plasma | mmol/L | 8.5% | 9.4% | 8% | 0.06 | 0.75 | RCPA QAP
Phosphate (Urine) | Urine | mmol/L | 26.4% | 26.5% | 15% | 3 | 20 | RCPA QAP
Porphobilinogen (Urine) | Urine | umol/L | 17% | 31% | 25% | 5 | 20 | RCPA QAP
Porphyrins (Blood) | Blood | umol/Lrc |  |  | 20% | 0.4 | 2 | RCPA QAP
Porphyrins (Faeces) | Faeces | umol/kg |  |  | 20% | 40 | 200 | RCPA QAP
Porphyrins (Plasma) | Plasma | nmol/L |  |  | 20% | 4 | 20 | RCPA QAP
Porphyrins (Urine) | Urine | nmol/L |  |  | 20% | 60 | 300 | RCPA QAP
Posaconazole | Serum/Plasma | mg/L |  |  | 10% | 0.1 | 1 | RCPA QAP
Potassium | Serum/Plasma | mmol/L | 4.2% | 4.4% | 5% | 0.2 | 4 | RCPA QAP
Potassium (Blood) | Blood | mmol/L | 4.2% | 4.4% | 5% | 0.2 | 4 | RCPA QAP
Potassium (Urine) | Urine | mmol/L | 27.1% | 23.2% | 10% | 2 | 20 | RCPA QAP
Prealbumin | Serum/Plasma | g/L | 10.9% | 19.1% | 20% | 0.04 | 0.2 | RCPA QAP
Procalcitonin | Serum/Plasma | ug/L |  |  | 30% | 0.05 | 0.15 | RCPA QAP
Procollagen type 1 N-terminal propeptide | Serum/Plasma | ug/L | 8.8% | 34.9% | 10% | 5 | 50 | RCPA QAP
Progesterone | Serum/Plasma | nmol/L |  |  | 15% | 2 | 10 | RCPA QAP
Prolactin | Serum/Plasma | mIU/L | 16% | 27% | 10% | 40 | 400 | RCPA QAP
Prostate Specific Antigen | Serum/Plasma | ug/L | 18.1% | 72.4% | 8% | 0.4 | 5 | RCPA QAP
Prostate Specific Antigen (Free) | Serum/Plasma | ug/L | 8.8% | 57.7% | 15% | 0.2 | 1.4 | RCPA QAP
Protein | Serum/Plasma | g/L | 2.6% | 4.6% | 5% | 3 | 60 | RCPA QAP
Protein (CSF) | CSF | g/L |  |  | 10% | 0.05 | 0.5 | RCPA QAP
Protein (Urine) | Urine | g/L | 39.6% | 17.8% | 10% | 0.1 | 1 | RCPA QAP
Prothrombin Time | Plasma | s | 4% | 6.8% | 20% | 2 | 10 | RCPA QAP
Pseudocholinesterase | Serum/Plasma | kU/L | 7% | 10.4% | 10% | 0.5 | 5 | RCPA QAP
Renin (conc) | Plasma | mU/L | 30.1% | 41.6% | 30% | 6 | 20 | RCPA QAP
Rheumatoid factor | Serum | IU/mL |  |  | 20% | 12 | 60 | RCPA QAP
Salicylate (mass) | Serum/Plasma | mg/L |  |  | 10% | 14 | 143 | RCPA QAP
Salicylate (molar) | Serum/Plasma | mmol/L |  |  | 10% | 0.1 | 1 | RCPA QAP
Sex Hormone Binding Globulin | Serum/Plasma | nmol/L | 9.7% | 35.6% | 12% | 6 | 50 | RCPA QAP
Sirolimus (Blood) | Blood | ug/L |  |  | 10% | 0.5 | 5 | RCPA QAP
Sodium | Serum/Plasma | mmol/L | 0.5% | 1.2% | 2% | 3 | 150 | RCPA QAP
Sodium (Blood) | Blood | mmol/L | 0.5% | 1.2% | 2% | 3 | 150 | RCPA QAP
Sodium (Sweat) | Sweat | mmol/L |  |  | 5% | 2 | 40 | RCPA QAP
Sodium (Urine) | Urine | mmol/L | 24% | 26.8% | 10% | 4 | 40 | RCPA QAP
Soluble FMS Like Tyrosine Kinase 1 | Serum | ng/L | 4% | 7.9% | 10% | 10 | 100 | RCPA QAP
Tacrolimus (Blood) | Blood | ug/L |  |  | 10% | 0.5 | 5 | RCPA QAP
Testosterone | Serum/Plasma | nmol/L | 12.4% | 17.7% | 15% | 0.4 | 2.7 | RCPA QAP
Theophylline (mass) | Serum/Plasma | mg/L |  |  | 10% | 0.5 | 5.4 | RCPA QAP
Theophylline (molar) | Serum/Plasma | umol/L |  |  | 10% | 3 | 30 | RCPA QAP
Thrombin Time (Plasma) | Plasma | s |  |  | 25% | 3 | 12 | RCPA QAP
Thyroglobulin | Serum/Plasma | ug/L | 12.8% | 29.2% | 12% | 0.2 | 1.7 | RCPA QAP
Thyroid Stimulating Hormone | Serum/Plasma | mU/L | 21.2% | 35% | 20% | 0.1 | 0.5 | RCPA QAP
Thyroxine (free) | Serum/Plasma | pmol/L | 7.7% | 12% | 12% | 1.5 | 12 | RCPA QAP
Thyroxine (total) | Serum/Plasma | nmol/L | 6.2% | 12.1% | 10% | 12 | 120 | RCPA QAP
Tobramycin | Serum/Plasma | mg/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Total Iron Binding Capacity | Serum/Plasma | umol/L |  |  | 8% | 4 | 50 | RCPA QAP
Transferrin | Serum/Plasma | g/L | 3% | 4.3% | 8% | 0.2 | 2.5 | RCPA QAP
Triglycerides | Serum/Plasma | mmol/L | 20.5% | 30% | 12% | 0.2 | 1.6 | RCPA QAP
Triiodothyronine (free) | Serum/Plasma | pmol/L | 6% | 16.5% | 20% | 0.7 | 3.5 | RCPA QAP
Triiodothyronine (total) | Serum/Plasma | nmol/L | 9.3% | 11.6% | 15% | 0.2 | 1.3 | RCPA QAP
Troponin I (hs) | Serum/Plasma | ng/L | 14% | 63% | 20% | 2 | 10 | RCPA QAP
Troponin I (POCT) | Blood | ug/L | 14% | 63% | 20% | 0.01 | 0.05 | RCPA QAP
Troponin T (hs) | Serum/Plasma | ng/L | 30% | 90% | 20% | 10 | 50 | RCPA QAP
Troponin T (POCT) | Blood | ng/L | 30% | 90% | 20% | 10 | 50 | RCPA QAP
Urate (mmol/L) | Serum/Plasma | mmol/L | 9% | 17.6% | 8% | 0.03 | 0.38 | RCPA QAP
Urate (umol/L) | Serum/Plasma | umol/L | 9% | 17.6% | 8% | 30 | 380 | RCPA QAP
Urate (Urine) | Urine | mmol/L | 24.7% | 22.1% | 15% | 0.3 | 2 | RCPA QAP
Urea | Serum/Plasma | mmol/L | 13.9% | 18.5% | 12% | 0.5 | 4 | RCPA QAP
Urea (Urine) | Urine | mmol/L | 22.7% | 25.9% | 10% | 20 | 200 | RCPA QAP
Valproate (mass) | Serum/Plasma | mg/L |  |  | 10% | 4 | 36 | RCPA QAP
Valproate (molar) | Serum/Plasma | umol/L |  |  | 10% | 25 | 250 | RCPA QAP
Vancomycin | Serum/Plasma | mg/L |  |  | 10% | 2 | 20.3 | RCPA QAP
Vitamin A (Serum) | Serum | umol/L | 13.6% | 19% | 12% | 0.2 | 1.7 | RCPA QAP
Vitamin B1 (Plasma) | Plasma | nmol/L | 4.8% | 12% | 10% | 15 | 150 | RCPA QAP
Vitamin B12 | Serum | pmol/L | 15% | 69% | 15% | 18 | 120 | RCPA QAP
Vitamin B2 | Blood | nmol/L | 5.8% | 10% | 25% | 9 | 36 | RCPA QAP
Vitamin B6 (Blood) | Blood | nmol/L | 20% | 34% | 15% | 9 | 60 | RCPA QAP
Vitamin B6 (Plasma) | Plasma | nmol/L |  |  | 15% | 9 | 60 | RCPA QAP
Vitamin C | Serum/Plasma | umol/L |  |  | 20% | 7 | 35 | RCPA QAP
Vitamin D (25-hydroxy) | Serum/Plasma | nmol/L | 6.9% | 26.1% | 15% | 9 | 60 | RCPA QAP
Voriconazole | Serum/Plasma | mg/L |  |  | 10% | 0.2 | 2 | RCPA QAP
Xanthochromia - Bilirubin | CSF | AU |  |  | 20% | 0.002 | 0.007 | RCPA QAP
Xanthochromia - Haemoglobin | CSF | AU |  |  | 20% | 0.02 | 0.1 | RCPA QAP
Zinc | Serum/Plasma | umol/L | 9.3% | 9.4% | 10% | 2 | 20 | RCPA QAP

## Tables
Analyte Lookup | 17-Hydroxyprogesterone | Quality Stringency | Yes No
| Acid Phosphatase (Prostatic) | Optimal | N
| Acid Phosphatase (Tartrate Resistant) | Desirable | Y
| Acid Phosphatase (Total) | Minimal
Activated Partial Thromboplastin Time
Adrenaline
Adrenaline (Urine)
Adrenocorticotropic hormone (mass)
| Adrenocorticotropic hormone (molar) | Matrix | Run
| Alanine aminotransferase | Serum/Plasma | 1
| Albumin | Plasma | 2
| Albumin (CSF) | Serum | 3
| Albumin (Urine) | Blood | 4
| Aldosterone | Urine
| Alkaline Phosphatase | CSF
| Alphafetoprotein | Pleural fluid
| Amikacin | Peritoneal fluid
| Aminolevulinic acid | Drain fluid
| Amiodarone | Dialysate
Amitriptyline
Ammonia
Amphetamine (class) (Urine)
Amylase
Amylase (Urine)
Androstenedione
Angiotensin Converting Enzyme
Ant Mullerian Hormone
Apolipoprotein A1
Apolipoprotein B
Aspartate aminotransferase
B-2-Microglobulin
Barbiturates (Urine)
Benzodiazepines (class) (Urine)
Beta Hydroxybutyrate (mmol/L)
Beta Hydroxybutyrate (umol/L)
Bicarbonate
Bile Acids (Total)
Bilirubin (Conjugated)
Bilirubin (CSF)
Bilirubin (Neonatal)
Bilirubin (Total)
Blood (Faeces)
Brain Natriuretic Peptide
C3 Complement
C4 Complement
CA 125
CA 15-3
CA 19-9
Caeruloplasmin
Calcitonin
Calcium
Calcium (Ionised)
Calcium (Urine)
Cannabinoids (class) (Urine)
Carbamazepine (mass)
Carboxyhaemoglobin
Carcinoembryonic Antigen
Chloride
Chloride (Blood)
Chloride (Sweat)
Chloride (Urine)
Cholesterol
Cholinesterase
Ciclosporin (Blood)
CO2 (Partial pressure)
Cocaine (class) (Urine)
Conductivity
Copper
Cortisol
Cortisol (Saliva)
Cortisol (Urine)
C-Peptide
C-Reactive Protein
C-Reactive Protein (high sensitivity)
Creatine Kinase
Creatine Kinase MB
Creatinine
Creatinine (Blood)
Creatinine (enzymatic)
Creatinine (Urine)
CTX Beta-Crosslaps
D-Dimer
Dehydroepiandrosterone Sulfate
Deoxypyridinoline crosslinks (Urine)
Digoxin (mass)
Dihydrotestosterone
Dopamine (Urine)
Ethanol
Everolimus
Ferritin
Fibrinogen
Folate
Folate (Red Cell)
Follicle Stimulating Hormone
Free Beta Human Chorionic Gonadotropin
Fructosamine
Gamma glutamyltransferase
Gastrin
Gentamicin
Glucose
Glucose (Blood)
Glucose (CSF)
Glucose (Urine)
Growth Hormone
Haematocrit (POCT)
Haemoglobin (Plasma Free)
Haemoglobin (POCT)
Haptoglobin
HbA1c (IFCC)
HbA1c (NGSP)
HDL Cholesterol
Holotranscobalamin
Homocysteine
Homovanillic Acid (Urine)
Human Chorionic Gonadotropin
Human Epididymal Protein 4
Hydroxyindole Acetic Acid
IGF Binding Protein 3
Immunoglobulin A
Immunoglobulin G
Immunoglobulin G (CSF)
Immunoglobulin M
Insulin
Insulin Like Growth Factor 1 (mass)
Insulin Like Growth Factor 1 (molar)
Iron
Itraconazole
Lactate
Lactate (Blood)
Lactate (CSF)
Lactate Dehydrogenase
LDL Cholesterol
Lipase
Lipoprotein (a) (mass)
Lithium
Luteinising Hormon
Magnesium
Magnesium (Urine)
Metanephrine (Plasma)
Metanephrine (Urine)
Methadone Metabolite
Methaemoglobin (Blood)
Methotrexate
Methoxytyramine (Plasma)
Methoxytyramine (Urine)
Mycophenolate
Neuron Specific Enolase
Normetanephrine (Plasma)
Normetanephrine (Urine)
Nortriptyline
N-telopeptide (Urine)
N-terminal proBNP
Oestradiol
Oestriol
Opiates (class) (Urine)
Osmolality
Osmolality (Urine)
Oxygen partial pressure (Blood)
Oxyhaemoglobin (Blood)
Paracetamol (mass)
Paracetamol (molar)
Parathyroid Hormone (mass)
Parathyroid Hormone (molar)
pH (Blood)
pH (Fluid)
Phenobarbitone (mass)
Phenobarbitone (molar)
Phenytoin (mass)
Phenytoin (molar)
Phosphate
Phosphate (Urine)
Porphobilinogen (Urine)
Porphyrins (Blood)
Porphyrins (Faeces)
Porphyrins (Plasma)
Porphyrins (Urine)
Posaconazole
Potassium
Potassium (Blood)
Potassium (Urine)
Prealbumin
Procalcitonin
Procollagen type 1 N-terminal propeptide
Progesterone
Prolactin
Prostate Specific Antigen
Prostate Specific Antigen (Free)
Protein
Protein (CSF)
Protein (Urine)
Prothrombin Time
Pseudocholinesterase
Renin (conc)
Rheumatoid factor
Salicylate (mass)
Salicylate (molar)
Sex Hormone Binding Globulin
Sirolimus (Blood)
Sodium
Sodium (Blood)
Sodium (Sweat)
Sodium (Urine)
Soluble FMS Like Tyrosine Kinase 1
Tacrolimus (Blood)
Testosterone
Theophylline (mass)
Theophylline (molar)
Thrombin Time (Plasma)
Thyroglobulin
Thyroid Stimulating Hormone
Thyroxine (free)
Thyroxine (total)
Tobramycin
Total Iron Binding Capacity
Transferrin
Triglycerides
Triiodothyronine (free)
Triiodothyronine (total)
Troponin I (hs)
Troponin I (POCT)
Troponin T (hs)
Troponin T (POCT)
Urate (mmol/L)
Urate (umol/L)
Urate (Urine)
Urea
Urea (Urine)
Valproate (mass)
Valproate (molar)
Vancomycin
Vitamin A (Serum)
Vitamin B1 (Plasma)
Vitamin B12
Vitamin B2
Vitamin B6 (Blood)
Vitamin B6 (Plasma)
Vitamin C
Vitamin D (25-hydroxy)
Voriconazole
Xanthochromia - Bilirubin
Xanthochromia - Haemoglobin
Zinc

## Layout
Attribute | Range Address | Description
aps-abs | J4 | Absolute APS specification
aps-rel | J6 | Relative APS specification
x-range | O61:P160 | Results from the reference method
y-range | R61:S160 | Results from the comparison (new) method
output-range | C167 | Top left cell for writing the results of the regression analysis
regression-method | paba | Default regression method. May be paba, deming, wdeming only.
confidence-interval-method | default | Method for determining confidence interval. May be default or bootstrap only.
use-calculated-error-ratio | TRUE | Only relevant for Deming or Weighted Deming regression. May be TRUE or FALSE
error-ratio | 1 | Error ratio to be used by Deming or Weighted Deming regression
bland-altman-range | N162:U175 | Range of cells over which the difference plot is placed
scatter-plot-range | N177:U190 | Range of cells over which the scatter plot is placed
output-labels | TRUE | Whether to write the regression results with labels. May by TRUE or FALSE
chart-data-range | AP61 | Top left cell for writing data used to construct the chart. This should be some place out of the way.
p-days-range | B295:B319 | Range of cells containing the days attribute. Do not include the column header
p-runs-range | C295:C319 | Range of cells containing the runs attribute. Do not include the column header
p-results-range | D295:G319 | Range of cells containing the results of the precision study. Each level must be in a separate column
p-output-range | AD294 | Top left cell for writing the results of the imprecision analysis.