## Comp 1
| Reference Method |  | Proposed Method
Sample ID | 1st rep | 2nd rep | 1st rep | 2nd rep | X | Y
2229705716 | 10583 | 10172 | 10737 | 10746 | 10377.5 | 10741.5
2229706564 | 3879 | 4233 | 3934 | 3763 | 4056 | 3848.5
2229804242 | 2674 | 2634 | 2679 | 2632 | 2654 | 2655.5
2229801511 | 4865 | 4629 | 4932 | 5449 | 4747 | 5190.5
2229805263 | 1471 | 1448 | 1546 | 1563 | 1459.5 | 1554.5
2229805226 | 6018 | 5742 | 5985 | 6065 | 5880 | 6025
2229804358 | 3891 | 3851 | 3922 | 3801 | 3871 | 3861.5
2229803634 | 2519 | 2403 | 2593 | 2474 | 2461 | 2533.5
2229804436 | 1867 | 1737 | 1880 | 1836 | 1802 | 1858
2229804060 | 1565 | 1650 | 1622 | 1650 | 1607.5 | 1636
2229803948 | 4267 | 4391 | 4282 | 4401 | 4329 | 4341.5
2229802716 | 7862 | 7961 | 8076 | 8294 | 7911.5 | 8185
2229806103 | 1831 | 1766 | 1813 | 1694 | 1798.5 | 1753.5
2229807634 | 6651 | 6357 | 6805 | 6781 | 6504 | 6793
2229903100 | 9130 | 9883 | 9275 | 8837 | 9506.5 | 9056
2229903874 | 4837 | 4725 | 4877 | 4572 | 4781 | 4724.5
2229905006 | 2059 | 2185 | 2146 | 2121 | 2122 | 2133.5
2230003011 | 17052 | 18666 | 16905 | 16248 | 17859 | 16576.5
2230003937 | 7365 | 6763 | 7259 | 6954 | 7064 | 7106.5
2230106239 | 940 | 987 | 996 | 977 | 963.5 | 986.5
2230106308 | 434 | 383 | 456 | 506 | 408.5 | 481
2230104994 | 5693 | 5553 | 5677 | 5529 | 5623 | 5603
2230104292 | 4987 | 4860 | 5071 | 5021 | 4923.5 | 5046
2230106691 | 2754 | 2736 | 2873 | 2726 | 2745 | 2799.5
2230204176 | 1725 | 1752 | 1775 | 1668 | 1738.5 | 1721.5
2231404848 | 7192 | 6922 | 7454 | 7641 | 7057 | 7547.5
2231406897 | 3024 | 2979 | 3058 | 2961 | 3001.5 | 3009.5
2231407330 | 903 | 888 | 954 | 1019 | 895.5 | 986.5
2231503109 | 11962 | 12349 | 11793 | 11631 | 12155.5 | 11712
2231504117 | 1295 | 1286 | 1347 | 1485 | 1290.5 | 1416
2231502835 | 948 | 938 | 1000 | 997 | 943 | 998.5
2231503694 | 11643 | 10982 | 11982 | 11415 | 11312.5 | 11698.5
2231504139 | 2687 | 2713 | 2777 | 2881 | 2700 | 2829
2231503832 | 3688 | 3434 | 3714 | 3839 | 3561 | 3776.5
2231503992 | 1756 | 1666 | 1873 | 1925 | 1711 | 1899
2231506001 | 989 | 1023 | 1062 | 1022 | 1006 | 1042
2231504468 | 1182 | 1179 | 1570 | 1662 | 1180.5 | 1616
2231602794 | 9098 | 8916 | 9212 | 9549 | 9007 | 9380.5
2231603495 | 4333 | 4543 | 4427 | 4238 | 4438 | 4332.5
2231503532 | 3486 | 3726 | 3516 | 3101 | 3606 | 3308.5
Passing-Bablok - Analyse It
Equation | Y = 57.23 + 0.9988 X
Parameter | Estimate | Bootstrap 95% CI
Intercept | 57.228 | -8.3781 | 135.5327
Slope | 0.99879 | 0.96285 | 1.029
Weighted Deming fit - Analyse It
Equation | Y = 80.39 + 0.9949 X
Variance ratio X/Y (λ) | 1
Parameter | Estimate | 95% CI |  | Jackknife SE
Intercept | 80.3878 | 44.8203 | 115.9554 | 17.5695
Slope | 0.99491 | 0.97516 | 1.0147 | 0.009756

## Comp 2
| x | y
1 | 159.1321 | 155.9807
2 | 122.4098 | 113.5862
3 | 54.372 | 55.9217
4 | 77.6091 | 83.5145
5 | 117.9104 | 126.4048
6 | 34.3441 | 32.2263
7 | 133.3301 | 147.8675
8 | 106.1254 | 107.7836
9 | 13.6207 | 15.3532
10 | 107.59 | 96.6047
11 | 123.0559 | 133.7471
12 | 135.9369 | 153.0884
13 | 98.9796 | 102.2063
14 | 144.1098 | 132.8954
15 | 81.8654 | 73.1662
16 | 175.537 | 166.0806
17 | 104.3353 | 103.3146
18 | 141.3146 | 135.8667
19 | 196.7359 | 201.3003
20 | 26.2213 | 26.1947
21 | 81.6479 | 74.9445
22 | 128.692 | 118.4202
23 | 78.0854 | 75.8779
24 | 99.8525 | 101.7123
25 | 199.5089 | 197.7452
26 | 36.4482 | 39.7671
27 | 52.1088 | 47.7577
28 | 149.9271 | 169.1029
29 | 212.1657 | 195.6134
30 | 183.8476 | 195.8145
31 | 105.4596 | 120.8788
32 | 206.2364 | 176.1376
33 | 175.4151 | 170.3832
34 | 138.2155 | 122.7169
35 | 37.1996 | 40.417
36 | 175.7178 | 191.7702
37 | 178.9389 | 172.5062
38 | 37.8356 | 37.0861
39 | 59.5202 | 58.4855
40 | 149.8621 | 169.5276
Passing-Bablok - mcr
| Estimate | SE | LCI | UCI
Intercept | -0.46582 | NaN | -5.8837 | 4.9161
Slope | 0.9941 | NaN | 0.93365 | 1.0845
Weighted Deming - mcr
| Estimate | SE | LCI | UCI
Intercept | 1.1479 | 1.3741 | -1.6337 | 3.9296
Slope | 0.98687 | 0.021352 | 0.94364 | 1.0301
WEIGHTED DEMING REGRESSION FIT:
Variance ratio | 0.8
| Estimate | SE | LCI | UCI
Intercept | 1.1705 | 1.3544 | 1.5714 | -3.9124
Slope | 0.98691 | 0.021221 | 0.94395 | 1.0299
Passing-Bablok fit - Analyse It
Equation | y = -0.4658 + 0.9941 x
Parameter | Estimate | Bootstrap 95% CI
Intercept | -0.46582 | -5.5135 | 4.3473
Slope | 0.9941 | 0.93608 | 1.0819
Weighted Deming fit
Equation | y = 1.148 + 0.9869 x
Variance ratio X/Y (λ) | 1
Parameter | Estimate | 95% CI |  | Jackknife SE
Intercept | 1.1479 | -1.6337 | 3.9296 | 1.3741
Slope | 0.98687 | 0.94365 | 1.0301 | 0.021352
Weighted Deming fit
Equation | y = 1.171 + 0.9869 x
Variance ratio X/Y (λ) | 0.8
Parameter | Estimate | 95% CI |  | Jackknife SE
Intercept | 1.1705 | -1.5714 | 3.9124 | 1.3544
Slope | 0.98691 | 0.94395 | 1.0299 | 0.021221

## 5x5
Imprecision Claim
|  | Level 1 | Level 2 | Level 3 | Level 4
| Mean | 106 | 463 | 106 | 463
| Repeatability CV | 1% | 1% | 1% | 1%
| Within lab CV | 2.6% | 2.6% | 2.6% | 2.6%
Day | Run | Level 1 | Level 2 | Level 3 | Level 4
Day 1 | Run 1 | 95.1 | 418.2 | 95.1 | 418.2
Day 1 | Run 1 | 95.2 | 415.2 | 95.2 | 415.2
Day 1 | Run 1 | 95.8 | 416.9 | 95.8 | 416.9
Day 1 | Run 1 | 95.4 | 416.3 | 95.4 | 416.3
Day 1 | Run 1 | 95.9 | 416.8 | 95.9 | 416.8
Day 2 | Run 1 | 98.2 | 421.3 | 98.2 | 421.3
Day 2 | Run 1 | 96.6 | 416.7 | 96.6 | 416.7
Day 2 | Run 1 | 96.9 | 422.5 | 96.9 | 422.5
Day 2 | Run 1 | 96.2 | 419.9 | 96.2 | 419.9
Day 2 | Run 1 | 91.6 | 405.7 | 91.6 | 405.7
Day 3 | Run 1 | 94 | 413.3 | 94 | 413.3
Day 3 | Run 1 | 92 | 411.7 | 92 | 411.7
Day 3 | Run 1 | 93.1 | 413.8 | 93.1 | 413.8
Day 3 | Run 1 | 92.7 | 412.3 | 92.7 | 412.3
Day 3 | Run 1 | 93.7 | 414.6 | 93.7 | 414.6
Day 4 | Run 1 | 91.4 | 401.3 | 91.4 | 401.3
Day 4 | Run 1 | 91.6 | 398.7 | 91.6 | 398.7
Day 4 | Run 1 | 92.5 | 399 | 92.5 | 399
Day 4 | Run 1 | 91.2 | 398.2 | 91.2 | 398.2
Day 4 | Run 1 | 93 | 402.4 | 93 | 402.4
Day 5 | Run 1 | 92.9 | 401.4 | 92.9 | 401.4
Day 5 | Run 1 | 92.1 | 402 | 92.1 | 402
Day 5 | Run 1 | 94.7 | 404.2 | 94.7 | 404.2
Day 5 | Run 1 | 89.2 | 399.7 | 89.2 | 399.7
Day 5 | Run 1 | 88 | 398.6 | 88 | 398.6

## 20x2x2
Day | Run | Level1 | EP05_A1 | Mean | 100.3899 | Simulation
Day 1 | Run 1 | 105.9125 | 242 | SST | 1019.5938 | Repeatability | 3%
Day 1 | Run 1 | 103.4594 | 246 | SSA | 248.5388 | Between run | 1%
Day 1 | Run 2 | 100.3746 | 245 | SSB | 68.1579 | Between day | 2%
Day 1 | Run 2 | 103.7659 | 246 | SSE | 400.5985 | Total | 3.742%
Day 2 | Run 1 | 101.5234 | 243 | SSAB | 302.2987
Day 2 | Run 1 | 101.3348 | 242 | DF T | 79
Day 2 | Run 2 | 102.7242 | 238 | DF A | 19
Day 2 | Run 2 | 99.3245 | 238 | DF B | 1 |  | Mean | Repeatability CV | Repeatability SD | Within Laboratory CV | Within Laboratory SD
Day 3 | Run 1 | 104.6394 | 247 | DF AB | 19 |  | 100.3899 | 3.152% | 3.1646 | 3.763% | 3.7774
Day 3 | Run 1 | 94.1409 | 239 | DF E | 40
Day 3 | Run 2 | 95.332 | 241 | MSA | 13.081
Day 3 | Run 2 | 98.0019 | 240 | MSB | 68.1579
Day 4 | Run 1 | 104.5949 | 249 | MSAB | 15.9105 | Component | % of Total | CV | Exact / MLS 95% CI |  | SD
Day 4 | Run 1 | 98.5307 | 241 | MSE | 10.015 | Repeatability | 70.19% | 3.152% | 0.025881 | 0.040334 | 3.1646
Day 4 | Run 2 | 96.0395 | 250 | F AB | 1.5887 | Between Run | 29.81% | 2.054% | 0% | 3.775% | 2.0625
Day 4 | Run 2 | 93.7346 | 245 | F A | 1.3061 | Within Day | 100% | 3.763% | 0.031651 | 0.049621 | 3.7774
Day 5 | Run 1 | 103.817 | 246 | F B | 6.8056 | Between Day | 0% | 0% | 0% | 2.019% | 0
Day 5 | Run 1 | 100.9089 | 242 | N | 80 | Within Laboratory | 1 | 0.037627 | 0.033627 | 0.045952 | 3.7774
Day 5 | Run 2 | 95.2471 | 243 | Num A | 20
Day 5 | Run 2 | 103.2794 | 240 | Num B | 2 | ANOVA
Day 6 | Run 1 | 98.5788 | 244 | Num E | 2
Day 6 | Run 1 | 100.6577 | 245 | Var A | 0 | Source | SS | DF | MS | Expected MS
Day 6 | Run 2 | 104.575 | 251 | Var AB | 2.9477 | Day | 248.5388 | 19 | 13.081 | σ2Error + 2σ2Run(Day) + 4σ2Day
Day 6 | Run 2 | 102.1076 | 247 | Var E | 10.015 | Run(Day) | 370.4565 | 20 | 18.5228 | σ2Error + 2σ2Run(Day)
Day 7 | Run 1 | 98.1142 | 241 | SD A | 0 | Error | 400.5985 | 40 | 10.015 | σ2Error
Day 7 | Run 1 | 104.7764 | 246 | SD AB | 1.7169
Day 7 | Run 2 | 102.4656 | 245 | SD E | 3.1646
Day 7 | Run 2 | 99.9385 | 247 | SD WL | 3.6004
Day 8 | Run 1 | 103.0619 | 245 | CV A | 0 | Result Variance Component Analysis:
Day 8 | Run 1 | 102.9668 | 245 | CV AB | 0.017102 | -----------------------------------
Day 8 | Run 2 | 90.7506 | 243 | CV E | 0.031524
Day 8 | Run 2 | 93.4009 | 245 | CV WL | 0.035864 | Name | DF | SS | MS | VC | SD | CV
Day 9 | Run 1 | 100.5039 | 243 | DF WL | 74.2632 | Total | 72.0038 |  |  | 14.2689 | 3.7774 | 3.7627
Day 9 | Run 1 | 99.8629 | 239 | SD WL LCL | 3.1029 | Day | 19 | 248.5388 | 13.081 | 0 | 0 | 0
Day 9 | Run 2 | 98.9348 | 244 | SD WL UCL | 4.2893 | Day:Run | 20 | 370.4565 | 18.5228 | 4.2539 | 2.0625 | 2.0545
Day 9 | Run 2 | 102.1839 | 245 | SD E LCL | 2.5982 | Error | 40 | 400.5985 | 10.015 | 10.015 | 3.1646 | 3.1524
Day 10 | Run 1 | 95.9554 | 244 | SD E UCL | 4.0492
Day 10 | Run 1 | 108.4718 | 246 | CV WL LCL | 0.030909 | Mean: 100.3899 (N = 80)
Day 10 | Run 2 | 99.4938 | 247 | CV WL UCL | 0.042726
Day 10 | Run 2 | 95.756 | 239 | CV E LCL | 0.025881
Day 11 | Run 1 | 106.6435 | 252 | CV E UCL | 0.040334
Day 11 | Run 1 | 98.8522 | 251 | CV E UCL | 0.047943
Day 11 | Run 2 | 99.4934 | 247
Day 11 | Run 2 | 98.6346 | 241
Day 12 | Run 1 | 99.2624 | 249
Day 12 | Run 1 | 99.0536 | 248
Day 12 | Run 2 | 106.0952 | 251
Day 12 | Run 2 | 105.2823 | 246
Day 13 | Run 1 | 100.6536 | 242 |  |  | Experimental Design: balanced ¦ Method: ANOVA ¦ * VC set to 0 ¦ adapted MS used for total DF
Day 13 | Run 1 | 105.1235 | 240
Day 13 | Run 2 | 94.837 | 251
Day 13 | Run 2 | 97.9927 | 245
Day 14 | Run 1 | 99.0028 | 246
Day 14 | Run 1 | 103.745 | 249
Day 14 | Run 2 | 103.0862 | 248
Day 14 | Run 2 | 104.1824 | 240
Day 15 | Run 1 | 100.3203 | 247
Day 15 | Run 1 | 100.6459 | 248
Day 15 | Run 2 | 100.1346 | 245
Day 15 | Run 2 | 93.8468 | 246
Day 16 | Run 1 | 99.5437 | 240
Day 16 | Run 1 | 95.5514 | 238
Day 16 | Run 2 | 98.5671 | 239
Day 16 | Run 2 | 94.7814 | 242
Day 17 | Run 1 | 102.0144 | 241
Day 17 | Run 1 | 100.1264 | 244
Day 17 | Run 2 | 102.4432 | 245
Day 17 | Run 2 | 97.4825 | 248
Day 18 | Run 1 | 103.3301 | 244
Day 18 | Run 1 | 100.1662 | 244
Day 18 | Run 2 | 102.3017 | 237
Day 18 | Run 2 | 96.3192 | 242
Day 19 | Run 1 | 96.5498 | 241
Day 19 | Run 1 | 99.8726 | 239
Day 19 | Run 2 | 102.8303 | 247
Day 19 | Run 2 | 100.0121 | 245
Day 20 | Run 1 | 106.0571 | 247
Day 20 | Run 1 | 104.19 | 240
Day 20 | Run 2 | 100.0701 | 245
Day 20 | Run 2 | 102.8511 | 242

## anova
| Day | Run | Level1 | Day | Run 1 | Run 1 | Run 2 | Run 2 |  | Mean Day | Devsq Day | n |  | Devsq Error | Devsq Error |  |  | Devsq total |  |  |  | Mean Day
1 | Day 1 | Run 1 | 105.9125 | Day 1 | 105.9125 | 103.4594 | 100.3746 | 103.7659 |  | 103.3781 | 8.9296 | 4 | 35.7184 | 3.0089 | 5.7504 | 2 | 2 | 30.4998 | 9.4222 | 0.00023314 | 11.3975
2 | Day 1 | Run 1 | 103.4594 | Day 2 | 101.5234 | 101.3348 | 102.7242 | 99.3245 |  | 101.2267 | 0.70029 | 4 | 2.8012 | 0.017783 | 5.7792 | 2 | 2 | 1.2848 | 0.89286 | 5.4491 | 1.1351
3 | Day 1 | Run 2 | 100.3746 | Day 3 | 104.6394 | 94.1409 | 95.332 | 98.0019 |  | 98.0285 | 5.5759 | 4 | 22.3037 | 55.1086 | 3.5642 | 2 | 2 | 18.0582 | 39.0493 | 25.5824 | 5.7026
4 | Day 1 | Run 2 | 103.7659 | Day 4 | 104.5949 | 98.5307 | 96.0395 | 93.7346 |  | 98.2249 | 4.6871 | 4 | 18.7485 | 18.387 | 2.6563 | 2 | 2 | 17.6819 | 3.4565 | 18.926 | 44.2934
5 | Day 2 | Run 1 | 101.5234 | Day 5 | 103.817 | 100.9089 | 95.2471 | 103.2794 |  | 100.8131 | 0.17913 | 4 | 0.7165 | 4.2284 | 32.259 | 2 | 2 | 11.745 | 0.26941 | 26.448 | 8.3495
6 | Day 2 | Run 1 | 101.3348 | Day 6 | 98.5788 | 100.6577 | 104.575 | 102.1076 |  | 101.4798 | 1.1879 | 4 | 4.7517 | 2.161 | 3.044 | 2 | 2 | 3.28 | 0.07174 | 17.5155 | 2.9507
7 | Day 2 | Run 2 | 102.7242 | Day 7 | 98.1142 | 104.7764 | 102.4656 | 99.9385 |  | 101.3237 | 0.87192 | 4 | 3.4877 | 22.192 | 3.193 | 2 | 2 | 5.1786 | 19.241 | 4.3084 | 0.20376
8 | Day 2 | Run 2 | 99.3245 | Day 8 | 103.0619 | 102.9668 | 90.7506 | 93.4009 |  | 97.545 | 8.0931 | 4 | 32.3725 | 0.0045143 | 3.512 | 2 | 2 | 7.1394 | 6.6407 | 92.9157 | 48.8463
9 | Day 3 | Run 1 | 104.6394 | Day 9 | 100.5039 | 99.8629 | 98.9348 | 102.1839 |  | 100.3714 | 0.00034209 | 4 | 0.0013684 | 0.20544 | 5.2784 | 2 | 2 | 0.013009 | 0.27768 | 2.1173 | 3.2185
10 | Day 3 | Run 1 | 94.1409 | Day 10 | 95.9554 | 108.4718 | 99.4938 | 95.756 |  | 99.9192 | 0.22152 | 4 | 0.88607 | 78.3301 | 6.9857 | 2 | 2 | 19.6647 | 65.3172 | 0.80298 | 21.4734
11 | Day 3 | Run 2 | 95.332 | Day 11 | 106.6435 | 98.8522 | 99.4934 | 98.6346 |  | 100.906 | 0.26632 | 4 | 1.0653 | 30.3522 | 0.36876 | 2 | 2 | 39.1078 | 2.3645 | 0.80363 | 3.0809
12 | Day 3 | Run 2 | 98.0019 | Day 12 | 99.2624 | 99.0536 | 106.0952 | 105.2823 |  | 102.4234 | 4.1351 | 4 | 16.5406 | 0.021779 | 0.33042 | 2 | 2 | 1.2713 | 1.7855 | 32.5511 | 23.936
13 | Day 4 | Run 1 | 104.5949 | Day 13 | 100.6536 | 105.1235 | 94.837 | 97.9927 |  | 99.6517 | 0.54487 | 4 | 2.1795 | 9.9899 | 4.9793 | 2 | 2 | 0.069562 | 22.4072 | 30.8342 | 5.7463
14 | Day 4 | Run 1 | 98.5307 | Day 14 | 99.0028 | 103.745 | 103.0862 | 104.1824 |  | 102.5041 | 4.47 | 4 | 17.8799 | 11.2444 | 0.6009 | 2 | 2 | 1.924 | 11.2571 | 7.27 | 14.3834
15 | Day 4 | Run 2 | 96.0395 | Day 15 | 100.3203 | 100.6459 | 100.1346 | 93.8468 |  | 98.7369 | 2.7323 | 4 | 10.9293 | 0.053018 | 19.7682 | 2 | 2 | 0.0048454 | 0.065548 | 0.065158 | 42.8116
16 | Day 4 | Run 2 | 93.7346 | Day 16 | 99.5437 | 95.5514 | 98.5671 | 94.7814 |  | 97.1109 | 10.7519 | 4 | 43.0075 | 7.9692 | 7.1656 | 2 | 2 | 0.71611 | 23.4112 | 3.3226 | 31.4548
17 | Day 5 | Run 1 | 103.817 | Day 17 | 102.0144 | 100.1264 | 102.4432 | 97.4825 |  | 100.5166 | 0.016061 | 4 | 0.064243 | 1.7822 | 12.3046 | 2 | 2 | 2.639 | 0.069416 | 4.2162 | 8.4532
18 | Day 5 | Run 1 | 100.9089 | Day 18 | 103.3301 | 100.1662 | 102.3017 | 96.3192 |  | 100.5293 | 0.01943 | 4 | 0.077718 | 5.0053 | 17.8954 | 2 | 2 | 8.6449 | 0.050057 | 3.655 | 16.5709
19 | Day 5 | Run 2 | 95.2471 | Day 19 | 96.5498 | 99.8726 | 102.8303 | 100.0121 |  | 99.8162 | 0.32908 | 4 | 1.3163 | 5.5204 | 3.9712 | 2 | 2 | 14.746 | 0.26759 | 5.9558 | 0.1427
20 | Day 5 | Run 2 | 103.2794 | Day 20 | 106.0571 | 104.19 | 100.0701 | 102.8511 |  | 103.2921 | 8.4227 | 4 | 33.691 | 1.7431 | 3.8669 | 2 | 2 | 32.1176 | 14.4409 | 0.10226 | 6.0575
21 | Day 6 | Run 1 | 98.5788
22 | Day 6 | Run 1 | 100.6577 | Mean Run | 101.3129 |  | 99.4669 |  |  |  |  | SSA | 248.5388 | SSE | 400.5985 |  | SST | 1019.5938
23 | Day 6 | Run 2 | 104.575 | Devsq Run | 0.85197 |  | 0.85197 |  |  |  |  | dfA | 19 | dfE | 40 |  | dfT | 79
24 | Day 6 | Run 2 | 102.1076 |  | 40 |  | 40 | SSB
25 | Day 7 | Run 1 | 98.1142 |  | 34.0789 |  | 34.0789 | 68.1579 | [5 empty] | Mean | 100.3899
26 | Day 7 | Run 1 | 104.7764 |  |  |  |  | dfB |  |  |  | df
27 | Day 7 | Run 2 | 102.4656 | N | 80 |  |  | 2 |  | SSAB | 302.2987
28 | Day 7 | Run 2 | 99.9385 | nDay | 20 |  |  |  |  | SSB(in A) | 370.4565
29 | Day 8 | Run 1 | 103.0619 | nRun | 2
30 | Day 8 | Run 1 | 102.9668 | nRep | 2
31 | Day 8 | Run 2 | 90.7506 |  | df | SS | MS | VC | SD | CV | k_i | num | den | n | c
32 | Day 8 | Run 2 | 93.4009 | Total | 2 df: 72.0038; 2 VC: 14.2689; 2 SD: 3.7774; 2 CV: 3.763%
33 | Day 9 | Run 1 | 100.5039 | Day | 19 | 248.5388 | 13.081 | 0 | 0 | 0% | 0.05 | 0 | 0 | 20 | 0.24675
34 | Day 9 | Run 1 | 99.8629 | Run (in day) | 20 | 370.4565 | 18.5228 | 4.2539 | 2.0625 | 2.054% | 0.047619 | 0.20257 | 0.0020517 | 2 | 0.012987
35 | Day 9 | Run 2 | 98.9348 | Error | 40 | 400.5985 | 10.015 | 10.015 | 3.1646 | 3.152% | 0.02439 | 0.24427 | 0.0014917 | 2 | 0.012987
36 | Day 9 | Run 2 | 102.1839
37 | Day 10 | Run 1 | 95.9554 |  | 2 k_i: num; 2 num: 0.19966
38 | Day 10 | Run 1 | 108.4718 |  | 2 k_i: den; 2 num: 0.0035434
39 | Day 10 | Run 2 | 99.4938 |  | 2 k_i: df_wl; 2 num: 56.3483
40 | Day 10 | Run 2 | 95.756
41 | Day 11 | Run 1 | 106.6435
42 | Day 11 | Run 1 | 98.8522
43 | Day 11 | Run 2 | 99.4934
44 | Day 11 | Run 2 | 98.6346
45 | Day 12 | Run 1 | 99.2624
46 | Day 12 | Run 1 | 99.0536
47 | Day 12 | Run 2 | 106.0952
48 | Day 12 | Run 2 | 105.2823
49 | Day 13 | Run 1 | 100.6536
50 | Day 13 | Run 1 | 105.1235
51 | Day 13 | Run 2 | 94.837
52 | Day 13 | Run 2 | 97.9927
53 | Day 14 | Run 1 | 99.0028
54 | Day 14 | Run 1 | 103.745
55 | Day 14 | Run 2 | 103.0862
56 | Day 14 | Run 2 | 104.1824
57 | Day 15 | Run 1 | 100.3203
58 | Day 15 | Run 1 | 100.6459
59 | Day 15 | Run 2 | 100.1346
60 | Day 15 | Run 2 | 93.8468
61 | Day 16 | Run 1 | 99.5437
62 | Day 16 | Run 1 | 95.5514
63 | Day 16 | Run 2 | 98.5671
64 | Day 16 | Run 2 | 94.7814
65 | Day 17 | Run 1 | 102.0144
66 | Day 17 | Run 1 | 100.1264
67 | Day 17 | Run 2 | 102.4432
68 | Day 17 | Run 2 | 97.4825
69 | Day 18 | Run 1 | 103.3301
70 | Day 18 | Run 1 | 100.1662
71 | Day 18 | Run 2 | 102.3017
72 | Day 18 | Run 2 | 96.3192
73 | Day 19 | Run 1 | 96.5498
74 | Day 19 | Run 1 | 99.8726
75 | Day 19 | Run 2 | 102.8303
76 | Day 19 | Run 2 | 100.0121
77 | Day 20 | Run 1 | 106.0571
78 | Day 20 | Run 1 | 104.19
79 | Day 20 | Run 2 | 100.0701
80 | Day 20 | Run 2 | 102.8511
Unbalanced Data
| Day | Run | Level1
1 | Day 1 | Run 1 | 105.9125
2 | Day 1 | Run 1 | 103.4594
3 | Day 1 | Run 2 | 100.3746
4 | Day 1 | Run 2 | 103.7659
5 | Day 2 | Run 1 | 101.5234
6 | Day 2 | Run 1 | 101.3348
7 | Day 2 | Run 2
8 | Day 2 | Run 2
9 | Day 3 | Run 1 | 104.6394
10 | Day 3 | Run 1 | 94.1409
11 | Day 3 | Run 2 | 95.332
12 | Day 3 | Run 2 | 98.0019
13 | Day 4 | Run 1 | 104.5949
14 | Day 4 | Run 1 | 98.5307
15 | Day 4 | Run 2 | 96.0395
16 | Day 4 | Run 2 | 93.7346
17 | Day 5 | Run 1 | 103.817
18 | Day 5 | Run 1 | 100.9089
19 | Day 5 | Run 2 | 95.2471
20 | Day 5 | Run 2 | 103.2794
21 | Day 6 | Run 1 | 98.5788
22 | Day 6 | Run 1 | 100.6577
23 | Day 6 | Run 2 | 104.575
24 | Day 6 | Run 2 | 102.1076
25 | Day 7 | Run 1 | 98.1142
26 | Day 7 | Run 1 | 104.7764
27 | Day 7 | Run 2 | 102.4656
28 | Day 7 | Run 2 | 99.9385
29 | Day 8 | Run 1 | 103.0619
30 | Day 8 | Run 1 | 102.9668
31 | Day 8 | Run 2 | 90.7506
32 | Day 8 | Run 2 | 93.4009
33 | Day 9 | Run 1 | 100.5039
34 | Day 9 | Run 1 | 99.8629
35 | Day 9 | Run 2 | 98.9348
36 | Day 9 | Run 2 | 102.1839
37 | Day 10 | Run 1 | 95.9554
38 | Day 10 | Run 1 | 108.4718
39 | Day 10 | Run 2 | 99.4938
40 | Day 10 | Run 2
41 | Day 11 | Run 1 | 106.6435
42 | Day 11 | Run 1 | 98.8522
43 | Day 11 | Run 2 | 99.4934
44 | Day 11 | Run 2 | 98.6346
45 | Day 12 | Run 1 | 99.2624
46 | Day 12 | Run 1 | 99.0536
47 | Day 12 | Run 2 | 106.0952
48 | Day 12 | Run 2 | 105.2823
49 | Day 13 | Run 1 | 100.6536
50 | Day 13 | Run 1 | 105.1235
51 | Day 13 | Run 2 | 94.837
52 | Day 13 | Run 2 | 97.9927
53 | Day 14 | Run 1 | 99.0028
54 | Day 14 | Run 1 | 103.745
55 | Day 14 | Run 2 | 103.0862
56 | Day 14 | Run 2 | 104.1824
57 | Day 15 | Run 1 | 100.3203
58 | Day 15 | Run 1 | 100.6459
59 | Day 15 | Run 2 | 100.1346
60 | Day 15 | Run 2 | 93.8468
61 | Day 16 | Run 1 | 99.5437
62 | Day 16 | Run 1 | 95.5514
63 | Day 16 | Run 2 | 98.5671
64 | Day 16 | Run 2 | 94.7814
65 | Day 17 | Run 1 | 102.0144
66 | Day 17 | Run 1 | 100.1264
67 | Day 17 | Run 2 | 102.4432
68 | Day 17 | Run 2 | 97.4825
69 | Day 18 | Run 1 | 103.3301
70 | Day 18 | Run 1 | 100.1662
71 | Day 18 | Run 2 | 102.3017
72 | Day 18 | Run 2 | 96.3192
73 | Day 19 | Run 1 | 96.5498
74 | Day 19 | Run 1 | 99.8726
75 | Day 19 | Run 2 | 102.8303
76 | Day 19 | Run 2 | 100.0121
77 | Day 20 | Run 1 | 106.0571
78 | Day 20 | Run 1 | 104.19
79 | Day 20 | Run 2 | 100.0701
80 | Day 20 | Run 2 | 102.8511

## Level1
MSA: Level1 | [6 empty] | v5.92
20x2x2 A1:C81
Filter: No filter
Last updated 16 September 2025 at 16:18 by Doug Chesher
Precision
|  | Mean | Repeatability CV | Repeatability SD | Within Laboratory CV | Within Laboratory SD
|  | 100.3899 | 3.152% | 3.1646 | 3.763% | 3.7774
| Component | % of Total | CV | Exact / MLS 95% CI |  | SD
| Repeatability | 70.19% | 3.152% | 2.588% | 4.033% | 3.1646
| Between Run | 29.81% | 2.054% | 0% | 3.775% | 2.0625
| Within Day | 100% | 3.763% | 3.165% | 4.962% | 3.7774
| Between Day | 0% | 0% | 0% | 2.019% | 0
| Within Laboratory | 100% | 3.763% | 3.363% | 4.595% | 3.7774
ANOVA
| Source | SS | DF | MS | Expected MS
| Day | 248.5388 | 19 | 13.081 | σ2Error + 2σ2Run(Day) + 4σ2Day
| Run(Day) | 370.4565 | 20 | 18.5228 | σ2Error + 2σ2Run(Day)
| Error | 400.5985 | 40 | 10.015 | σ2Error

## Layout
Attribute | Range Address | Description
import-aps | TRUE | Import APS values from spreadsheet
aps-abs | J4 | Absolute APS specification or cell address
aps-rel | J6 | Relative APS specification or cell address
x-range | O61:P160 | Results from the reference method
y-range | R61:S160 | Results from the comparison (new) method
output-range | C167 | Top left cell for writing the results of the regression analysis
regression-method | paba | Default regression method. May be paba, deming, wdeming only.
confidence-interval-method | default | Method for determining confidence interval. May be default or bootstrap only.
use-calculated-error-ratio | FALSE | Only relevant for Deming or Weighted Deming regression. May be TRUE or FALSE
error-ratio | 1 | Error ratio to be used by Deming or Weighted Deming regression
bland-altman-range | N162:U175 | Range of cells over which the difference plot is placed
scatter-plot-range | N177:U190 | Range of cells over which the scatter plot is placed
output-labels | TRUE | Whether to write the regression results with labels. May by TRUE or FALSE
chart-data-range | AP61 | Top left cell for writing data used to construct the chart. This should be some place out of the way.
p-days-range | B295:B319 | Range of cells containing the days attribute. Do not include the column header
p-runs-range | C295:C319 | Range of cells containing the runs attribute. Do not include the column header
p-results-range | D295:G319 | Range of cells containing the results of the precision study. Each level must be in a separate column
p-output-range | AD294 | Top left cell for writing the results of the imprecision analysis.
concordance-output | C176 | Top left cell for writing the results of the concordance analysis
qx-range | E61:E160 | Qualitative results from the reference method
qy-range | H61:H160 | Qualitative results from the comparison (new) method
q-output-range | C176 | Top left cell for writing the results of the concordance analysis

## Raw
Day | Run | Level1 | EP05_A1
Day 1 | Run 1 | 105.9125 | 242
Day 1 | Run 1 | 103.4594 | 246
Day 1 | Run 2 | 100.3746 | 245
Day 1 | Run 2 | 103.7659 | 246
Day 2 | Run 1 | 101.5234 | 243
Day 2 | Run 1 | 101.3348 | 242
Day 2 | Run 2 | 102.7242 | 238
Day 2 | Run 2 | 99.3245 | 238
Day 3 | Run 1 | 104.6394 | 247
Day 3 | Run 1 | 94.1409 | 239
Day 3 | Run 2 | 95.332 | 241
Day 3 | Run 2 | 98.0019 | 240
Day 4 | Run 1 | 104.5949 | 249
Day 4 | Run 1 | 98.5307 | 241
Day 4 | Run 2 | 96.0395 | 250
Day 4 | Run 2 | 93.7346 | 245
Day 5 | Run 1 | 103.817 | 246
Day 5 | Run 1 | 100.9089 | 242
Day 5 | Run 2 | 95.2471 | 243
Day 5 | Run 2 | 103.2794 | 240
Day 6 | Run 1 | 98.5788 | 244
Day 6 | Run 1 | 100.6577 | 245
Day 6 | Run 2 | 104.575 | 251
Day 6 | Run 2 | 102.1076 | 247
Day 7 | Run 1 | 98.1142 | 241
Day 7 | Run 1 | 104.7764 | 246
Day 7 | Run 2 | 102.4656 | 245
Day 7 | Run 2 | 99.9385 | 247
Day 8 | Run 1 | 103.0619 | 245
Day 8 | Run 1 | 102.9668 | 245
Day 8 | Run 2 | 90.7506 | 243
Day 8 | Run 2 | 93.4009 | 245
Day 9 | Run 1 | 100.5039 | 243
Day 9 | Run 1 | 99.8629 | 239
Day 9 | Run 2 | 98.9348 | 244
Day 9 | Run 2 | 102.1839 | 245
Day 10 | Run 1 | 95.9554 | 244
Day 10 | Run 1 | 108.4718 | 246
Day 10 | Run 2 | 99.4938 | 247
Day 10 | Run 2 | 95.756 | 239
Day 11 | Run 1 | 106.6435 | 252
Day 11 | Run 1 | 98.8522 | 251
Day 11 | Run 2 | 99.4934 | 247
Day 11 | Run 2 | 98.6346 | 241
Day 12 | Run 1 | 99.2624 | 249
Day 12 | Run 1 | 99.0536 | 248
Day 12 | Run 2 | 106.0952 | 251
Day 12 | Run 2 | 105.2823 | 246
Day 13 | Run 1 | 100.6536 | 242
Day 13 | Run 1 | 105.1235 | 240
Day 13 | Run 2 | 94.837 | 251
Day 13 | Run 2 | 97.9927 | 245
Day 14 | Run 1 | 99.0028 | 246
Day 14 | Run 1 | 103.745 | 249
Day 14 | Run 2 | 103.0862 | 248
Day 14 | Run 2 | 104.1824 | 240
Day 15 | Run 1 | 100.3203 | 247
Day 15 | Run 1 | 100.6459 | 248
Day 15 | Run 2 | 100.1346 | 245
Day 15 | Run 2 | 93.8468 | 246
Day 16 | Run 1 | 99.5437 | 240
Day 16 | Run 1 | 95.5514 | 238
Day 16 | Run 2 | 98.5671 | 239
Day 16 | Run 2 | 94.7814 | 242
Day 17 | Run 1 | 102.0144 | 241
Day 17 | Run 1 | 100.1264 | 244
Day 17 | Run 2 | 102.4432 | 245
Day 17 | Run 2 | 97.4825 | 248
Day 18 | Run 1 | 103.3301 | 244
Day 18 | Run 1 | 100.1662 | 244
Day 18 | Run 2 | 102.3017 | 237
Day 18 | Run 2 | 96.3192 | 242
Day 19 | Run 1 | 96.5498 | 241
Day 19 | Run 1 | 99.8726 | 239
Day 19 | Run 2 | 102.8303 | 247
Day 19 | Run 2 | 100.0121 | 245
Day 20 | Run 1 | 106.0571 | 247
Day 20 | Run 1 | 104.19 | 240
Day 20 | Run 2 | 100.0701 | 245
Day 20 | Run 2 | 102.8511 | 242

## 10x3x2
Days | Runs | Level1
Day 1 | Run 1 | 242
Day 1 | Run 1 | 246
Day 1 | Run 1 | 245
Day 1 | Run 1 | 246
Day 1 | Run 2 | 243
Day 1 | Run 2 | 242
Day 1 | Run 2 | 238
Day 1 | Run 2 | 238
Day 1 | Run 3 | 247
Day 1 | Run 3 | 239
Day 1 | Run 3 | 241
Day 1 | Run 3 | 240
Day 2 | Run 1 | 249
Day 2 | Run 1 | 241
Day 2 | Run 1 | 250
Day 2 | Run 1 | 245
Day 2 | Run 2 | 246
Day 2 | Run 2 | 242
Day 2 | Run 2 | 243
Day 2 | Run 2 | 240
Day 2 | Run 3 | 244
Day 2 | Run 3 | 245
Day 2 | Run 3 | 251
Day 2 | Run 3 | 247
Day 3 | Run 1 | 241
Day 3 | Run 1 | 246
Day 3 | Run 1 | 245
Day 3 | Run 1 | 247
Day 3 | Run 2 | 245
Day 3 | Run 2 | 245
Day 3 | Run 2 | 243
Day 3 | Run 2 | 245
Day 3 | Run 3 | 243
Day 3 | Run 3 | 239
Day 3 | Run 3 | 244
Day 3 | Run 3 | 245
Day 4 | Run 1 | 244
Day 4 | Run 1 | 246
Day 4 | Run 1 | 247
Day 4 | Run 1 | 239
Day 4 | Run 2 | 252
Day 4 | Run 2 | 251
Day 4 | Run 2 | 247
Day 4 | Run 2 | 241
Day 4 | Run 3 | 249
Day 4 | Run 3 | 248
Day 4 | Run 3 | 251
Day 4 | Run 3 | 246
Day 5 | Run 1 | 242
Day 5 | Run 1 | 240
Day 5 | Run 1 | 251
Day 5 | Run 1 | 245
Day 5 | Run 2 | 246
Day 5 | Run 2 | 249
Day 5 | Run 2 | 248
Day 5 | Run 2 | 240
Day 5 | Run 3 | 247
Day 5 | Run 3 | 248
Day 5 | Run 3 | 245
Day 5 | Run 3 | 246

## Quant
SUMMARY
| Assay |  |  |  |  | Units |  |  | APS Absolute
| Soluble FMS Like Tyrosine Kinase 1 |  |  |  |  | ng/L |  |  | 10
| Matrix | [7 empty] | APS Relative
| Serum | [7 empty] | 10%
| Instrument/Device | [7 empty] | APS Cutoff
| Some manufacturer device | [7 empty] | 100
Method principle
Chemiluminescence
IFU Insert Version
Evaluation Period
12/11/2023
Evaluation Completed By
Comments and Assessment
Decision
| Approved By | [6 empty] | Approval Date
Commencement Date
COMARISON USING PATIENT SAMPLES
Reference Method Instrument/Device
Reference Method Principle
Reference Method IFU Version
Diagnostic Cutoff for Concordance Assessment (Enter up to 4)
|  |  | Reference Thresholds |  |  | Proposed Thresholds
|  |  | 2000 |  |  | 2010
|  |  | 4000 |  |  | 4010
|  |  | 6000 |  |  | 6010
|  |  | Reference Method |  |  | Proposed Method
Reagent Lot No.
Calibrator Lot No.
|  |  | Reference Method |  |  | Proposed Method | [6 empty] | Reference Method |  |  | Proposed Method
| Sample ID |  | Operator | 1st rep | 2nd rep | Operator | 1st rep | 2nd rep | Inclusion | Comment |  |  | X R1 | X R2 | Mean X | Y R1 | Y R2 | Mean Y
| 1000 |  |  | 10583 | 10172 | JS | 10737 | 10746 | INCLUDE |  |  |  | 10583 | 10172 | 10377.5 | 10737 | 10746 | 10741.5 | [7 empty] | Mean | Relative Difference | X | Y
| 1001 |  |  | 3879 | 4233 | JS | 3934 | 3763 | INCLUDE |  |  |  | 3879 | 4233 | 4056 | 3934 | 3763 | 3848.5 | [7 empty] | 10559.5 | 0.034471 | 10377.5 | 10741.5 | 444.75 | 0.02956 | 444.75 | 0.16132 | 444.75 | -0.1022 | 100 | 0.1 | 100 | -0.1
| 1002 |  |  | 2674 | 2634 | JS | 2679 | 2632 | INCLUDE |  |  |  | 2674 | 2634 | 2654 | 2679 | 2632 | 2655.5 | [7 empty] | 3952.25 | -0.052502 | 4056 | 3848.5 | 17217.75 | 0.02956 | 17217.75 | 0.16132 | 17217.75 | -0.1022 | 17217.75 | 0.1 | 17217.75 | -0.1
| 1003 |  |  | 4865 | 4629 | JS | 4932 | 5449 | INCLUDE |  |  |  | 4865 | 4629 | 4747 | 4932 | 5449 | 5190.5 | [7 empty] | 2654.75 | 0.00056502 | 2654 | 2655.5
| 1004 |  |  | 1471 | 1448 | JS | 1546 | 1563 | INCLUDE |  |  |  | 1471 | 1448 | 1459.5 | 1546 | 1563 | 1554.5 | [7 empty] | 4968.75 | 0.089258 | 4747 | 5190.5
| 1005 |  |  | 6018 | 5742 | JS | 5985 | 6065 | INCLUDE |  |  |  | 6018 | 5742 | 5880 | 5985 | 6065 | 6025 | [7 empty] | 1507 | 0.063039 | 1459.5 | 1554.5
| 1006 |  |  | 3891 | 3851 | JS | 3922 | 3801 | INCLUDE |  |  |  | 3891 | 3851 | 3871 | 3922 | 3801 | 3861.5 | [7 empty] | 5952.5 | 0.02436 | 5880 | 6025
| 1007 |  |  | 2519 | 2403 | JS | 2593 | 2474 | INCLUDE |  |  |  | 2519 | 2403 | 2461 | 2593 | 2474 | 2533.5 | [7 empty] | 3866.25 | -0.0024572 | 3871 | 3861.5
| 1008 |  |  | 1867 | 1737 | JS | 1880 | 1836 | INCLUDE |  |  |  | 1867 | 1737 | 1802 | 1880 | 1836 | 1858 | [7 empty] | 2497.25 | 0.029032 | 2461 | 2533.5
| 1009 |  |  | 1565 | 1650 | JS | 1622 | 1650 | INCLUDE |  |  |  | 1565 | 1650 | 1607.5 | 1622 | 1650 | 1636 | [7 empty] | 1830 | 0.030601 | 1802 | 1858
| 1010 |  |  | 4267 | 4391 | JS | 4282 | 4401 | INCLUDE |  |  |  | 4267 | 4391 | 4329 | 4282 | 4401 | 4341.5 | [7 empty] | 1621.75 | 0.017574 | 1607.5 | 1636 |  |  | 408.5 | 408.5 | 408.5 | 465.2345 | 100 | 110 | 100 | 90
| 1011 |  |  | 7862 | 7961 | JS | 8076 | 8294 | INCLUDE |  |  |  | 7862 | 7961 | 7911.5 | 8076 | 8294 | 8185 | [7 empty] | 4335.25 | 0.0028833 | 4329 | 4341.5 |  |  | 17859 | 17859 | 17859 | 17894.6493 | 17859 | 19644.9 | 17859 | 16073.1
| 1012 |  |  | 1831 | 1766 | JS | 1813 | 1694 | INCLUDE |  |  |  | 1831 | 1766 | 1798.5 | 1813 | 1694 | 1753.5 | [7 empty] | 8048.25 | 0.033983 | 7911.5 | 8185
| 1013 |  |  | 6651 | 6357 | JS | 6805 | 6781 | INCLUDE |  |  |  | 6651 | 6357 | 6504 | 6805 | 6781 | 6793 | [7 empty] | 1776 | -0.025338 | 1798.5 | 1753.5
| 1014 |  |  | 9130 | 9883 | JS | 9275 | 8837 | INCLUDE |  |  |  | 9130 | 9883 | 9506.5 | 9275 | 8837 | 9056 | [7 empty] | 6648.5 | 0.043468 | 6504 | 6793
| 1015 |  |  | 4837 | 4725 | JS | 4877 | 4572 | INCLUDE |  |  |  | 4837 | 4725 | 4781 | 4877 | 4572 | 4724.5 | [7 empty] | 9281.25 | -0.048539 | 9506.5 | 9056
| 1016 |  |  | 2059 | 2185 | JS | 2146 | 2121 | INCLUDE |  |  |  | 2059 | 2185 | 2122 | 2146 | 2121 | 2133.5 | [7 empty] | 4752.75 | -0.011888 | 4781 | 4724.5
| 1017 |  |  | 17052 | 18666 | JS | 16905 | 16248 | INCLUDE |  |  |  | 17052 | 18666 | 17859 | 16905 | 16248 | 16576.5 | [7 empty] | 2127.75 | 0.0054048 | 2122 | 2133.5
| 1018 |  |  | 7365 | 6763 | JS | 7259 | 6954 | INCLUDE |  |  |  | 7365 | 6763 | 7064 | 7259 | 6954 | 7106.5 | [7 empty] | 17217.75 | -0.074487 | 17859 | 16576.5
| 1019 |  |  | 940 | 987 | JS | 996 | 977 | INCLUDE |  |  |  | 940 | 987 | 963.5 | 996 | 977 | 986.5 | [7 empty] | 7085.25 | 0.0059984 | 7064 | 7106.5
| 1020 |  |  | 434 | 383 | JS | 456 | 506 | INCLUDE |  |  |  | 434 | 383 | 408.5 | 456 | 506 | 481 | [7 empty] | 975 | 0.02359 | 963.5 | 986.5
| 1021 |  |  | 5693 | 5553 | JS | 5677 | 5529 | INCLUDE |  |  |  | 5693 | 5553 | 5623 | 5677 | 5529 | 5603 | [7 empty] | 444.75 | 0.16301 | 408.5 | 481
| 1022 |  |  | 4987 | 4860 | JS | 5071 | 5021 | INCLUDE |  |  |  | 4987 | 4860 | 4923.5 | 5071 | 5021 | 5046 | [7 empty] | 5613 | -0.0035632 | 5623 | 5603
| 1023 |  |  | 2754 | 2736 | JS | 2873 | 2726 | INCLUDE |  |  |  | 2754 | 2736 | 2745 | 2873 | 2726 | 2799.5 | [7 empty] | 4984.75 | 0.024575 | 4923.5 | 5046
| 1024 |  |  | 1725 | 1752 | JS | 1775 | 1668 | INCLUDE |  |  |  | 1725 | 1752 | 1738.5 | 1775 | 1668 | 1721.5 | [7 empty] | 2772.25 | 0.019659 | 2745 | 2799.5
| 1025 |  |  | 7192 | 6922 | JS | 7454 | 7641 | INCLUDE |  |  |  | 7192 | 6922 | 7057 | 7454 | 7641 | 7547.5 | [7 empty] | 1730 | -0.0098266 | 1738.5 | 1721.5
| 1026 |  |  | 3024 | 2979 | JS | 3058 | 2961 | INCLUDE |  |  |  | 3024 | 2979 | 3001.5 | 3058 | 2961 | 3009.5 | [7 empty] | 7302.25 | 0.067171 | 7057 | 7547.5
| 1027 |  |  | 903 | 888 | JS | 954 | 1019 | INCLUDE |  |  |  | 903 | 888 | 895.5 | 954 | 1019 | 986.5 | [7 empty] | 3005.5 | 0.0026618 | 3001.5 | 3009.5
| 1028 |  |  | 11962 | 12349 | JS | 11793 | 11631 | INCLUDE |  |  |  | 11962 | 12349 | 12155.5 | 11793 | 11631 | 11712 | [7 empty] | 941 | 0.096706 | 895.5 | 986.5
| 1029 |  |  | 1295 | 1286 | JS | 1347 | 1485 | INCLUDE |  |  |  | 1295 | 1286 | 1290.5 | 1347 | 1485 | 1416 | [7 empty] | 11933.75 | -0.037164 | 12155.5 | 11712
| 1030 |  |  | 948 | 938 | JS | 1000 | 997 | INCLUDE |  |  |  | 948 | 938 | 943 | 1000 | 997 | 998.5 | [7 empty] | 1353.25 | 0.09274 | 1290.5 | 1416
| 1031 |  |  | 11643 | 10982 | JS | 11982 | 11415 | INCLUDE |  |  |  | 11643 | 10982 | 11312.5 | 11982 | 11415 | 11698.5 | [7 empty] | 970.75 | 0.057172 | 943 | 998.5
| 1032 |  |  | 2687 | 2713 | JS | 2777 | 2881 | INCLUDE |  |  |  | 2687 | 2713 | 2700 | 2777 | 2881 | 2829 | [7 empty] | 11505.5 | 0.033549 | 11312.5 | 11698.5
| 1033 |  |  | 3688 | 3434 | JS | 3714 | 3839 | INCLUDE |  |  |  | 3688 | 3434 | 3561 | 3714 | 3839 | 3776.5 | [7 empty] | 2764.5 | 0.046663 | 2700 | 2829
| 1034 |  |  | 1756 | 1666 | JS | 1873 | 1925 | INCLUDE |  |  |  | 1756 | 1666 | 1711 | 1873 | 1925 | 1899 | [7 empty] | 3668.75 | 0.058739 | 3561 | 3776.5
| 1035 |  |  | 989 | 1023 | JS | 1062 | 1022 | INCLUDE |  |  |  | 989 | 1023 | 1006 | 1062 | 1022 | 1042 | [7 empty] | 1805 | 0.10416 | 1711 | 1899
| 1036 |  |  | 1182 | 1179 | JS | 1570 | 1662 | INCLUDE |  |  |  | 1182 | 1179 | 1180.5 | 1570 | 1662 | 1616 | [7 empty] | 1024 | 0.035156 | 1006 | 1042
| 1037 |  |  | 9098 | 8916 | JS | 9212 | 9549 | INCLUDE |  |  |  | 9098 | 8916 | 9007 | 9212 | 9549 | 9380.5 | [7 empty] | 1398.25 | 0.31146 | 1180.5 | 1616
| 1038 |  |  | 4333 | 4543 | JS | 4427 | 4238 | INCLUDE |  |  |  | 4333 | 4543 | 4438 | 4427 | 4238 | 4332.5 | [7 empty] | 9193.75 | 0.040625 | 9007 | 9380.5
| 1039 |  |  | 3486 | 3726 | JS | 3516 | 3101 | INCLUDE |  |  |  | 3486 | 3726 | 3606 | 3516 | 3101 | 3308.5 | [7 empty] | 4385.25 | -0.024058 | 4438 | 4332.5
[9 empty] | EXCLUDE | [16 empty] | 3457.25 | -0.086051 | 3606 | 3308.5
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
|  |  | x | y |  |  |  | Pearson Correlation Coefficient (R) | 0.99716
|  | N | 40 | 40 |  |  |  | Coefficient of Determination (R^2) | 0.99432
|  | Mean | 4526.45 | 4569.1125
|  | Passing-Bablock Regression |  |  | Non-parametric CI
|  |  | Coefficents | LCL | UCL
|  | Slope | 0.99879 | 0.97106 | 1.0277
|  | Intercept | 57.2281 | -8.7746 | 125.4356
Concordance
n x n concordance table
|  |  | <2000 | ≥2000-<4000 | ≥4000-<6000 | ≥6000-<10000 | ≥10000 | Total
|  | <2000 | 13 | 0 | 0 | 0 | 0 | 13
|  | ≥2000-<4000 | 0 | 9 | 0 | 0 | 0 | 9
|  | ≥4000-<6000 | 0 | 1 | 6 | 1 | 0 | 8
|  | ≥6000-<10000 | 0 | 0 | 0 | 6 | 0 | 6
|  | ≥10000 | 0 | 0 | 0 | 0 | 4 | 4
|  | Total | 13 | 10 | 6 | 7 | 4 | 40
|  | Overall Agreement | 0.95
|  | Expected Agreement | 0.22812
|  | Kappa | 0.93522 | Good agreement
|  | SE(Kappa) | 0.044335
|  | LCL | 0.84833
|  | UCL | 1.0221
TRUENESS ASSESSED USING STANDARD MATERIALS
Trueness verification based on RCPAQAP material
EQA Program Name
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
| ENTER below MIMIMUM of 6 and MAXIMUM of 8 levels in single or replicate values for test method | [15 empty] | Calculations
[17 empty] | Trueness using RCPAQAP materials |  |  |  | Cut-off Values in reportable units
| Sample ID | Global Target | Method Group Median | First Replicate | Second Replicate | Mean | Trueness
[17 empty] | sd (sx) | s2x | Bias | %Bias | High cut off | Low cut off |  | n
| 23-08 | 200 | 195 | 190 | 190 | 190 | Pass | [9 empty] | 0 | 0 | -5 | -2.5641 | 186.7269 | 203.2731 |  | 2
| 23-09 | 250 | 230 | 232 | 232 | 232 | Pass | [9 empty] | 0 | 0 | 2 | 0.86957 | 220.2419 | 239.7581 |  | 2
| 23-10 | 10 | 20 | 15 |  | 15 | Fail | [9 empty] | 0 | 0 | -5 | -25 | 18.8 | 21.2 |  | 1
| 23-11 | 35 | 38 | 35 | 35 | 35 | Fail | [9 empty] | 0 | 0 | -3 | -7.8947 | 36.3878 | 39.6122 |  | 2
| 23-12 | 90 | 91 | 89 | 89 | 89 | Pass | [9 empty] | 0 | 0 | -2 | -2.1978 | 87.1392 | 94.8608 |  | 2
| 23-13 | 75 | 70 | 68 | 67 | 67.5 | Pass | [9 empty] | 0.70711 | 0.26726 | -2.5 | -3.5714 | 67.0302 | 72.9698 |  | 2
| 23-14 | 100 | 100 | 95 | 95 | 95 | Fail | [9 empty] | 0 | 0 | -5 | -5 | 95.7574 | 104.2426 |  | 2
| 23-15 | 200 | 205 | [13 empty] | 0 | 0 | [5 empty] | 0
Method Group Median CV from EQA Program
3%
| Number of labs in EQA Method Group |  |  |  |  | Std Err of method (sa) | 0.013416
| 5 |  |  |  |  | Number of samples tested | 7
[6 empty] | Number of replicates | 2
Trueness verification based on other reference materials
Trueness is verified if the TEST method gives results that fall within the calculated lower and upper verification limits.
This form of verification should be used if (1) the two methods are substantially different e.g. the test method uses a different
analytical principle to the reference method, or (2) the manufacturer has not provided relevant bias information, or (3) an
EQA program does not exist that can be used for trueness verification, or (4) alternative reference materials suitable for
trueness verification are available e.g. calibrator materials from different batch numbers. The default minimum acceptable
imprecision is as published by RCPAQAP.
Name of Material Tested
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
ENTER below MIMIMUM of 3 different batches of calibrators
| ID | Target Value | First Replicate | Second Replicate | Third Replicate | Mean | Strict | t-test
[17 empty] | SE | Abs Bias | Bias CI Lower | Bias CI Upper | Max Bias | Low cut off | High cut off
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
| Allowable bias |  | Alpha
5%
PRECISION
Material Used
Material Lot No.
Reagent Lot No.
Calibrator Lot No.
Claimed Imprecision
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Mean | 106 | 463 | 106 | 463
|  | Repeatability CV | 1% | 1% | 1% | 1%
|  | Within lab CV | 2.6% | 2.6% | 2.6% | 2.6%
ENTER up to 5 values per day for each level
| Day | Run | Level 1 | Level 2 | Level 3 | Level 4 | [14 empty] | Mean | 93.56 | 409.628 | 93.56 | 409.628
| Day 1 | Run 1 | 95.1 | 418.2 | 95.1 | 418.2 | [14 empty] | SS Total | 144.58 | 1653.3904 | 144.58 | 1653.3904
| Day 1 | Run 1 | 95.2 | 415.2 | 95.2 | 415.2 | [14 empty] | SST | 83.752 | 1426.5864 | 83.752 | 1426.5864
| Day 1 | Run 1 | 95.8 | 416.9 | 95.8 | 416.9 | [14 empty] | SSE | 60.828 | 226.804 | 60.828 | 226.804
| Day 1 | Run 1 | 95.4 | 416.3 | 95.4 | 416.3 | [14 empty] | DF Total | 24 | 24 | 24 | 24
| Day 1 | Run 1 | 95.9 | 416.8 | 95.9 | 416.8 | [14 empty] | DF Error | 20 | 20 | 20 | 20
| Day 2 | Run 1 | 98.2 | 421.3 | 98.2 | 421.3 | [14 empty] | DF T | 4 | 4 | 4 | 4
| Day 2 | Run 1 | 96.6 | 416.7 | 96.6 | 416.7 | [14 empty] | MST | 20.938 | 356.6466 | 20.938 | 356.6466
| Day 2 | Run 1 | 96.9 | 422.5 | 96.9 | 422.5 | [14 empty] | MSE | 3.0414 | 11.3402 | 3.0414 | 11.3402
| Day 2 | Run 1 | 96.2 | 419.9 | 96.2 | 419.9 | [14 empty] | F | 6.8843 | 31.4498 | 6.8843 | 31.4498
| Day 2 | Run 1 | 91.6 | 405.7 | 91.6 | 405.7 | [14 empty] | N | 25 | 25 | 25 | 25
| Day 3 | Run 1 | 94 | 413.3 | 94 | 413.3 | [14 empty] | P | 5 | 5 | 5 | 5
| Day 3 | Run 1 | 92 | 411.7 | 92 | 411.7 | [14 empty] | Sum GrpN2 | 125 | 125 | 125 | 125
| Day 3 | Run 1 | 93.1 | 413.8 | 93.1 | 413.8 | [14 empty] | Var Error | 3.0414 | 11.3402 | 3.0414 | 11.3402
| Day 3 | Run 1 | 92.7 | 412.3 | 92.7 | 412.3 | [14 empty] | Var B | 3.5793 | 69.0613 | 3.5793 | 69.0613
| Day 3 | Run 1 | 93.7 | 414.6 | 93.7 | 414.6 | [14 empty] | SD Error | 1.744 | 3.3675 | 1.744 | 3.3675
| Day 4 | Run 1 | 91.4 | 401.3 | 91.4 | 401.3 | [14 empty] | CV Error | 0.01864 | 0.0082209 | 0.01864 | 0.0082209
| Day 4 | Run 1 | 91.6 | 398.7 | 91.6 | 398.7 | [14 empty] | SD B | 1.8919 | 8.3103 | 1.8919 | 8.3103
| Day 4 | Run 1 | 92.5 | 399 | 92.5 | 399 | [14 empty] | CV B | 0.020221 | 0.020287 | 0.020221 | 0.020287
| Day 4 | Run 1 | 91.2 | 398.2 | 91.2 | 398.2 | [14 empty] | SD WL | 2.5731 | 8.9667 | 2.5731 | 8.9667
| Day 4 | Run 1 | 93 | 402.4 | 93 | 402.4 | [14 empty] | CV WL | 0.027502 | 0.02189 | 0.027502 | 0.02189
| Day 5 | Run 1 | 92.9 | 401.4 | 92.9 | 401.4 | [14 empty] | DF WL | 9.3662 | 5.0658 | 9.3662 | 5.0658
| Day 5 | Run 1 | 92.1 | 402 | 92.1 | 402 | [14 empty] | ChiSq E | 36.7605 | 36.7605 | 36.7605 | 36.7605
| Day 5 | Run 1 | 94.7 | 404.2 | 94.7 | 404.2 | [14 empty] | ChiSq WL | 21.5956 | 14.658 | 21.5956 | 14.658
| Day 5 | Run 1 | 89.2 | 399.7 | 89.2 | 399.7 | [14 empty] | F Error | 1.3557 | 1.3557 | 1.3557 | 1.3557
| Day 5 | Run 1 | 88 | 398.6 | 88 | 398.6 | [14 empty] | F WL | 1.5185 | 1.701 | 1.5185 | 1.701
|  | Repeatability | 1.3206 | 1.8351 | 1.3206 | 1.8351
|  | Between day | 1.8919 | 8.3103 | 1.8919 | 8.3103
|  | Within lab | 2.5731 | 8.9667 | 2.5731 | 8.9667
|  | Repeatability CV | 1.864% | 0.8221% | 1.864% | 0.8221%
|  | Between day CV | 2.022% | 92.68% | 73.53% | 92.68%
|  | Within lab CV | 2.75% | 2.189% | 2.75% | 2.189%
Upper Verification Limits
|  | Repeatability UVL | 1.356% | 1.356% | 1.356% | 1.356%
|  | Within lab UVL | 3.948% | 4.423% | 3.948% | 4.423%
Comparison to Claim
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Repeatability | Fail | Pass | Fail | Pass
|  | Within Lab Precision | Pass | Pass | Pass | Pass
Assay Capability
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Mean | 2.5731 | 8.9667 | 2.5731 | 8.9667
|  | RCPA APS | 388.6% | 111.5% | 388.6% | 111.5%
|  | Capability (σ) | 141.3139 | 50.9478 | 141.3139 | 50.9478
|  | Acceptability | Good | Good | Good | Good
Comparison to Biological Variation APS
|  | CVi | 4% |  |  | Specification
|  | CVg | 7.9% |  | Imprecision | 2%
|  | Required Category | Desirable |  | Bias | 2.214%
|  | alpha | 0.05 |  | Total Error | 5.503%
Upper Verification Limits
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Repeatability | 2.711% | 2.711% | 2.711% | 2.711%
|  | Within Lab Precision | 3.037% | 3.402% | 3.037% | 3.402%
Imprecision Verification Result
|  | Repeatability | Pass | Pass | Pass | Pass
|  | Within Lab Precision | Pass | Pass | Pass | Pass
REFERENCE INTERVALS
Reference Intervals that will be used
Reference Interval Verification
| LRL | URL |  |  |  | LRL | URL
| 30 | 110 |  |  |  | 30 | 110
| Sample ID | Operator | Result |  |  | Sample ID | Operator | Result
1 | Test 1 | T1 | 61 | 0 |  | Test 1 | T1 | 61 | 0
2 | Test 2 | T2 | 97 | 0 |  | Test 2 | T2 | 97 | 0
3 | Test 3 | T3 | 93 | 0 |  | Test 3 | T3 | 93 | 0
4 | Test 4 | T4 | 48 | 0 |  | Test 4 | T4 | 48 | 0
5 | Test 5 | T5 | 52 | 0 |  | Test 5 | T5 | 52 | 0
6 | Test 6 | T6 | 52 | 0 |  | Test 6 | T6 | 52 | 0
7 | Test 7 | T7 | 41 | 0 |  | Test 7 | T7 | 41 | 0
8 | Test 8 | T8 | 75 | 0 |  | Test 8 | T8 | 75 | 0
9 | Test 9 | T9 | 123 | 1 |  | Test 9 | T9 | 123 | 1
10 | Test 10 | T10 | 55 | 0 |  | Test 10 | T10 | 55 | 0
11 | Test 11 | T11 | 56 | 0 |  | Test 11 | T11 | 56 | 0
12 | Test 12 | T12 | 111 | 1 |  | Test 12 | T12 | 111 | 1
13 | Test 13 | T13 | 74 | 0 |  | Test 13 | T13 | 74 | 0
14 | Test 14 | T14 | 73 | 0 |  | Test 14 | T14 | 73 | 0
15 | Test 15 | T15 | 67 | 0 |  | Test 15 | T15 | 67 | 0
16 | Test 16 | T16 | 95 | 0 |  | Test 16 | T16 | 95 | 0
17 | Test 17 | T17 | 55 | 0 |  | Test 17 | T17 | 55 | 0
18 | Test 18 | T18 | 44 | 0 |  | Test 18 | T18 | 44 | 0
19 | Test 19 | T19 | 59 | 0 |  | Test 19 | T19 | 59 | 0
20 | Test 20 | T20 | 112 | 1 |  | Test 20 | T20 | 112 | 1
21 | [5 empty] | Test 21 | T21 | 95 | 0
22 | [5 empty] | Test 22 | T22 | 90 | 0
23 | [5 empty] | Test 23 | T23 | 44 | 0
24 | [5 empty] | Test 24 | T24 | 65 | 0
25 | [5 empty] | Test 25 | T25 | 62 | 0
26 | [5 empty] | Test 26 | T26 | 46 | 0
27 | [5 empty] | Test 27 | T27 | 97 | 0
28 | [5 empty] | Test 28 | T28 | 56 | 0
29 | [5 empty] | Test 29 | T29 | 70 | 0
30 | [5 empty] | Test 30 | T30 | 71 | 0
31 | [5 empty] | Test 31 | T31 | 48 | 0
32 | [5 empty] | Test 32 | T32 | 64 | 0
33 | [5 empty] | Test 33 | T33 | 44 | 0
34 | [5 empty] | Test 34 | T34 | 75 | 0
35 | [5 empty] | Test 35 | T35 | 57 | 0
36 | [5 empty] | Test 36 | T36 | 52 | 0
37 | [5 empty] | Test 37 | T37 | 69 | 0
38 | [5 empty] | Test 38 | T38 | 95 | 0
39 | [5 empty] | Test 39 | T39 | 51 | 0
40 | [5 empty] | Test 40 | T40 | 74 | 0
|  | Reference interval verification | Fail |  |  |  | Reference interval verification | Pass
|  | N | 20 |  |  |  | N | 40
|  | Outliers | 3 |  |  |  | Outliers | 3
|  | p | 0.015902 |  |  |  | p | 0.13815
LIMIT OF BLANK, DETECTION, QUANTITATION
| Date of analysis |  |  | Reagent Lot No.
| Operator ID |  |  | Calibrator Lot No.
Limit of Blank (LoB)
| ENTER the LoB claim |  | 0.9 | ng/L
ENTER min of 2 replicates, of min 2 blank samples, of min 3 days, for min of 20 values
[9 empty] | Bounds for observed proportion of results relative to a detection capability claim | [10 empty] | Values ≤ LoB claim
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Total number of measurements (n) |  |  | % Observed proportion boundary | [7 empty] | 1 | 1 | 1
|  | Day 1 | 2022-10-23 | 0.5 | 0.8 |  |  |  | 20 |  |  | 85% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.5 | 0.5 |  |  |  | 30 |  |  | 86.67% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.6 | 0.6 |  |  |  | 40 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 2 | 2022-10-24 | 0.7 | 0.7 |  |  |  | 50 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.7 | 0.7 |  |  |  | 60 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.8 | 0.9 |  |  |  | 70 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 3 | 2022-10-25 | 0.5 | 0.6 |  |  |  | 80 |  |  | 91.25% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.4 | 0.5 |  |  |  | 90 |  |  | 91.11% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.6 | 0.7 |  |  |  | 100 |  |  | 91% |  |  |  |  | n <= LoB claim |  | 27
[9 empty] | 150 |  |  | 92% |  |  |  |  | % n<= LoB claim |  | 150 | %
[9 empty] | 200 |  |  | 92.5%
| Total number of values |  | 18 |  | Min Req | 15 |  |  | 250 |  |  | 92.8%
| mean of blank |  | 0.62778 |  | % Min Req | 83.33% |  |  | 300 |  |  | 93%
| SD of blank |  | 0.12826 | [5 empty] | 400 |  |  | 93.25%
Limit of blank (LoB) |  |  | 0.84232 | 0 | Alpha | 0.05 |  |  | 500 |  |  | 93.4%
|  |  | Pass |  | C_beta | 1.6727 |  |  | 1000 |  |  | 93.8%
Limit of Detection (LoD)
| ENTER the LoD claim |  | 1.1 | ng/L
Values ≥ LoB claim
|  | ENTER min of 2 replicates, of min 2 LoD claim samples, of min 3 days, for min of 20 values | [17 empty] | 1 | 1
1
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  | Selected samples should have a concentration at the claimed LoD | [11 empty] | 1 | 1
|  | Day 1 | 2022-10-23 | 1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 0.9 | 0.8 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 1 | 1.1 | [14 empty] | 1 | 1
|  | Day 2 | 2022-10-24 | 1.1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 0.9 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 1.1 | 1.1 | [14 empty] | 1 | 1
|  | Day 3 | 2022-10-25 | 1.2 | 1.1 | [11 empty] | n >=LoB claim |  | 17
|  |  | 2022-10-25 | 1.1 | 0.9 | [11 empty] | % n >= LoB claim |  | 94.4444 | %
|  |  | 2022-10-25 | 1 | 1.1
|  |  | Sample SD | 0.094281 | 0.099381
|  |  |  | 0.0088889 | 0.0098765 | #N/A
| Total number of values |  | 18 |  | Min Req | 15
| Pooled SD |  | 0.096864 |  | % Min Req | 83.33%
|  | Limit of Detection | 1.062 | 0
|  |  | Pass |  | Alpha | 0.05
[5 empty] | C_beta | 1.6727
Limit of Quantitaion (LoQ)
| Enter the LoQ claim |  | 1.5 | ng/L
ENTER min of 2 replicates, of min 2 LoQ claim samples, of min 3 days, for min of 20 values
% difference from means
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Sample 1 | Sample 2 | Sample 3 |  |  |  |  | Values < TEa |  |  |  | % difference calculations
|  | Day 1 | 2022-10-23 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-23 | 13 | 14 |  |  |  | 14.5985 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 14.5985 | 8.0292 | #VALUE!
|  |  | 2022-10-23 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  | Day 2 | 2022-10-24 | 16 | 15 |  |  |  | 5.1095 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 5.1095 | 1.4599 | #VALUE!
|  |  | 2022-10-24 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-24 | 17 | 14 |  |  |  | 11.6788 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 11.6788 | 8.0292 | #VALUE!
|  | Day 3 | 2022-10-25 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-25 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-25 | 18 | 16 |  |  |  | 18.2482 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 18.2482 | 5.1095 | #VALUE!
|  | Sample mean |  | 15.2222 | 15.2222 | [7 empty] | n >= TEa |  | 0
|  |  | Sample SD | 1.4741 | 0.78567 | [7 empty] | % n >= TEa |  | 0
|  |  | Sample CV | 9.684% | 5.161%
| Total number of values |  | 18 |  | Min Req | 15
| Total number of samples |  | 2 |  | % Min Req | 83.33%
Grand Mean |  |  | 15.2222
[5 empty] | Alpha | 0.05
| TEa |  | #REF! | % | C_beta | 1.6727
Fail
Note on B380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on C380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit
Note on G380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on H380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit

## Qual
SUMMARY
| Assay |  |  |  |  | Units |  |  | APS Absolute
| Soluble FMS Like Tyrosine Kinase 1 |  |  |  |  | #REF! |  |  | #REF!
| Matrix | [7 empty] | APS Relative
| Serum | [7 empty] | #REF!
| Instrument/Device | [7 empty] | APS Cutoff
| Some manufacturer device | [7 empty] | #REF!
Method principle
Chemiluminescence
IFU Insert Version
Evaluation Period
12/11/2023
Evaluation Completed By
Comments and Assessment
Decision
| Approved By | [6 empty] | Approval Date
Commencement Date
COMARISON USING PATIENT SAMPLES
Reference Method Instrument/Device
Reference Method Principle
Reference Method IFU Version
Diagnostic Cutoff for Concordance Assessment (Enter up to 4)
|  |  | Reference Thresholds |  |  | Proposed Thresholds
|  |  | 2000 |  |  | 2010
|  |  | 4000 |  |  | 4010
|  |  | 6000 |  |  | 6010
|  |  | Reference Method |  |  | Proposed Method
Reagent Lot No.
Calibrator Lot No.
|  |  | Reference Method |  |  | Proposed Method | [6 empty] | Reference Method |  |  | Proposed Method
| Sample ID |  | Operator | 1st rep | 2nd rep | Operator | 1st rep | 2nd rep | Inclusion | Comment |  |  | X R1 | X R2 | Mean X | Y R1 | Y R2 | Mean Y
| 1000 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | Mean | Relative Difference | X | Y
| 1001 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 10559.5 | 0.034471 | 10377.5 | 10741.5 | 444.75 | 0.02956 | 444.75 | 0.16132 | 444.75 | -0.1022 | 100 | 0.1 | 100 | -0.1
| 1002 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 3952.25 | -0.052502 | 4056 | 3848.5 | 17217.75 | 0.02956 | 17217.75 | 0.16132 | 17217.75 | -0.1022 | 17217.75 | 0.1 | 17217.75 | -0.1
| 1003 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 2654.75 | 0.00056502 | 2654 | 2655.5
| 1004 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 4968.75 | 0.089258 | 4747 | 5190.5
| 1005 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 1507 | 0.063039 | 1459.5 | 1554.5
| 1006 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 5952.5 | 0.02436 | 5880 | 6025
| 1007 |  |  | A |  |  | A |  | INCLUDE |  |  |  | A |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 3866.25 | -0.0024572 | 3871 | 3861.5
| 1008 |  |  | A |  |  | B |  | INCLUDE |  |  |  | A |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 2497.25 | 0.029032 | 2461 | 2533.5
| 1009 |  |  | A |  |  | B |  | INCLUDE |  |  |  | A |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 1830 | 0.030601 | 1802 | 1858
| 1010 |  |  | B |  |  | A |  | INCLUDE |  |  |  | B |  | #DIV/0! | A |  | #DIV/0! | [7 empty] | 1621.75 | 0.017574 | 1607.5 | 1636 |  |  | 408.5 | 408.5 | 408.5 | 465.2345 | 100 | 110 | 100 | 90
| 1011 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 4335.25 | 0.0028833 | 4329 | 4341.5 |  |  | 17859 | 17859 | 17859 | 17894.6493 | 17859 | 19644.9 | 17859 | 16073.1
| 1012 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 8048.25 | 0.033983 | 7911.5 | 8185
| 1013 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 1776 | -0.025338 | 1798.5 | 1753.5
| 1014 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 6648.5 | 0.043468 | 6504 | 6793
| 1015 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 9281.25 | -0.048539 | 9506.5 | 9056
| 1016 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 4752.75 | -0.011888 | 4781 | 4724.5
| 1017 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 2127.75 | 0.0054048 | 2122 | 2133.5
| 1018 |  |  | B |  |  | B |  | INCLUDE |  |  |  | B |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 17217.75 | -0.074487 | 17859 | 16576.5
| 1019 |  |  | B |  |  | C |  | INCLUDE |  |  |  | B |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 7085.25 | 0.0059984 | 7064 | 7106.5
| 1020 |  |  | C |  |  | B |  | INCLUDE |  |  |  | C |  | #DIV/0! | B |  | #DIV/0! | [7 empty] | 975 | 0.02359 | 963.5 | 986.5
| 1021 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 444.75 | 0.16301 | 408.5 | 481
| 1022 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 5613 | -0.0035632 | 5623 | 5603
| 1023 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 4984.75 | 0.024575 | 4923.5 | 5046
| 1024 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 2772.25 | 0.019659 | 2745 | 2799.5
| 1025 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 1730 | -0.0098266 | 1738.5 | 1721.5
| 1026 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 7302.25 | 0.067171 | 7057 | 7547.5
| 1027 |  |  | C |  |  | C |  | INCLUDE |  |  |  | C |  | #DIV/0! | C |  | #DIV/0! | [7 empty] | 3005.5 | 0.0026618 | 3001.5 | 3009.5
| 1028 |  |  | C |  |  | D |  | INCLUDE |  |  |  | C |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 941 | 0.096706 | 895.5 | 986.5
| 1029 |  |  | C |  |  | D |  | INCLUDE |  |  |  | C |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 11933.75 | -0.037164 | 12155.5 | 11712
| 1030 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 1353.25 | 0.09274 | 1290.5 | 1416
| 1031 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 970.75 | 0.057172 | 943 | 998.5
| 1032 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 11505.5 | 0.033549 | 11312.5 | 11698.5
| 1033 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 2764.5 | 0.046663 | 2700 | 2829
| 1034 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 3668.75 | 0.058739 | 3561 | 3776.5
| 1035 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 1805 | 0.10416 | 1711 | 1899
| 1036 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 1024 | 0.035156 | 1006 | 1042
| 1037 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 1398.25 | 0.31146 | 1180.5 | 1616
| 1038 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 9193.75 | 0.040625 | 9007 | 9380.5
| 1039 |  |  | D |  |  | D |  | INCLUDE |  |  |  | D |  | #DIV/0! | D |  | #DIV/0! | [7 empty] | 4385.25 | -0.024058 | 4438 | 4332.5
[9 empty] | EXCLUDE | [16 empty] | 3457.25 | -0.086051 | 3606 | 3308.5
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
EXCLUDE
|  |  | x | y |  |  |  | Pearson Correlation Coefficient (R)
|  | N | 0 | 0 |  |  |  | Coefficient of Determination (R^2)
Mean
|  | Passing-Bablock Regression |  |  | Bootstrap CI
|  |  | Coefficents | LCL | UCL
|  | Slope | 0.99879 | 0.96203 | 1.0288
|  | Intercept | 57.2281 | -8.8973 | 126.2927
Concordance
n x n concordance table
|  |  | A | B | C | D | Total
|  | A | 8 | 2 | 0 | 0 | 10
|  | B | 1 | 8 | 1 | 0 | 10
|  | C | 0 | 1 | 7 | 2 | 10
|  | D | 0 | 0 | 0 | 10 | 10
|  | Total | 9 | 11 | 8 | 12 | 40
|  | Overall Agreement | 0.825
|  | Expected Agreement | 0.25
|  | Kappa | 0.76667 | Fair agreement
|  | SE(Kappa) | 0.079736
|  | LCL | 0.61039
|  | UCL | 0.92295
TRUENESS ASSESSED USING STANDARD MATERIALS
Trueness verification based on RCPAQAP material
EQA Program Name
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
| ENTER below MIMIMUM of 6 and MAXIMUM of 8 levels in single or replicate values for test method | [15 empty] | Calculations
[17 empty] | Trueness using RCPAQAP materials |  |  |  | Cut-off Values in reportable units
| Sample ID | Global Target | Method Group Median | First Replicate | Second Replicate | Mean | Trueness
[17 empty] | sd (sx) | s2x | Bias | %Bias | High cut off | Low cut off |  | n
| 23-08 | 200 | 195 | 190 | 190 | 190 | Pass | [9 empty] | 0 | 0 | -5 | -2.5641 | 186.7269 | 203.2731 |  | 2
| 23-09 | 250 | 230 | 232 | 232 | 232 | Pass | [9 empty] | 0 | 0 | 2 | 0.86957 | 220.2419 | 239.7581 |  | 2
| 23-10 | 10 | 20 | 15 |  | 15 | Fail | [9 empty] | 0 | 0 | -5 | -25 | 18.8 | 21.2 |  | 1
| 23-11 | 35 | 38 | 35 | 35 | 35 | Fail | [9 empty] | 0 | 0 | -3 | -7.8947 | 36.3878 | 39.6122 |  | 2
| 23-12 | 90 | 91 | 89 | 89 | 89 | Pass | [9 empty] | 0 | 0 | -2 | -2.1978 | 87.1392 | 94.8608 |  | 2
| 23-13 | 75 | 70 | 68 | 67 | 67.5 | Pass | [9 empty] | 0.70711 | 0.26726 | -2.5 | -3.5714 | 67.0302 | 72.9698 |  | 2
| 23-14 | 100 | 100 | 95 | 95 | 95 | Fail | [9 empty] | 0 | 0 | -5 | -5 | 95.7574 | 104.2426 |  | 2
| 23-15 | 200 | 205 | [13 empty] | 0 | 0 | [5 empty] | 0
Method Group Median CV from EQA Program
3%
| Number of labs in EQA Method Group |  |  |  |  | Std Err of method (sa) | 0.013416
| 5 |  |  |  |  | Number of samples tested | 7
[6 empty] | Number of replicates | 2
Trueness verification based on other reference materials
Trueness is verified if the TEST method gives results that fall within the calculated lower and upper verification limits.
This form of verification should be used if (1) the two methods are substantially different e.g. the test method uses a different
analytical principle to the reference method, or (2) the manufacturer has not provided relevant bias information, or (3) an
EQA program does not exist that can be used for trueness verification, or (4) alternative reference materials suitable for
trueness verification are available e.g. calibrator materials from different batch numbers. The default minimum acceptable
imprecision is as published by RCPAQAP.
Name of Material Tested
Date of Analysis
Operator ID
Reagent Lot No
Calibrator Lot No
ENTER below MIMIMUM of 3 different batches of calibrators
| ID | Target Value | First Replicate | Second Replicate | Third Replicate | Mean | Strict | t-test
[17 empty] | SE | Abs Bias | Bias CI Lower | Bias CI Upper | Max Bias | Low cut off | High cut off
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
[17 empty] | 0 |  | #VALUE! | #VALUE! | 0 | 0 | 0
| Allowable bias |  | Alpha
5%
PRECISION
Material Used
Material Lot No.
Reagent Lot No.
Calibrator Lot No.
Claimed Imprecision
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Mean | 106 | 463 | 106 | 463
|  | Repeatability CV | 1% | 1% | 1% | 1%
|  | Within lab CV | 2.6% | 2.6% | 2.6% | 2.6%
ENTER up to 5 values per day for each level
| Day | Run | Level 1 | Level 2 | Level 3 | Level 4 | [14 empty] | Mean | 93.56 | 409.628 | 93.56 | 409.628
| Day 1 | Run 1 | 95.1 | 418.2 | 95.1 | 418.2 | [14 empty] | SS Total | 144.58 | 1653.3904 | 144.58 | 1653.3904
| Day 1 | Run 1 | 95.2 | 415.2 | 95.2 | 415.2 | [14 empty] | SST | 83.752 | 1426.5864 | 83.752 | 1426.5864
| Day 1 | Run 1 | 95.8 | 416.9 | 95.8 | 416.9 | [14 empty] | SSE | 60.828 | 226.804 | 60.828 | 226.804
| Day 1 | Run 1 | 95.4 | 416.3 | 95.4 | 416.3 | [14 empty] | DF Total | 24 | 24 | 24 | 24
| Day 1 | Run 1 | 95.9 | 416.8 | 95.9 | 416.8 | [14 empty] | DF Error | 20 | 20 | 20 | 20
| Day 2 | Run 1 | 98.2 | 421.3 | 98.2 | 421.3 | [14 empty] | DF T | 4 | 4 | 4 | 4
| Day 2 | Run 1 | 96.6 | 416.7 | 96.6 | 416.7 | [14 empty] | MST | 20.938 | 356.6466 | 20.938 | 356.6466
| Day 2 | Run 1 | 96.9 | 422.5 | 96.9 | 422.5 | [14 empty] | MSE | 3.0414 | 11.3402 | 3.0414 | 11.3402
| Day 2 | Run 1 | 96.2 | 419.9 | 96.2 | 419.9 | [14 empty] | F | 6.8843 | 31.4498 | 6.8843 | 31.4498
| Day 2 | Run 1 | 91.6 | 405.7 | 91.6 | 405.7 | [14 empty] | N | 25 | 25 | 25 | 25
| Day 3 | Run 1 | 94 | 413.3 | 94 | 413.3 | [14 empty] | P | 5 | 5 | 5 | 5
| Day 3 | Run 1 | 92 | 411.7 | 92 | 411.7 | [14 empty] | Sum GrpN2 | 125 | 125 | 125 | 125
| Day 3 | Run 1 | 93.1 | 413.8 | 93.1 | 413.8 | [14 empty] | Var Error | 3.0414 | 11.3402 | 3.0414 | 11.3402
| Day 3 | Run 1 | 92.7 | 412.3 | 92.7 | 412.3 | [14 empty] | Var B | 3.5793 | 69.0613 | 3.5793 | 69.0613
| Day 3 | Run 1 | 93.7 | 414.6 | 93.7 | 414.6 | [14 empty] | SD Error | 1.744 | 3.3675 | 1.744 | 3.3675
| Day 4 | Run 1 | 91.4 | 401.3 | 91.4 | 401.3 | [14 empty] | CV Error | 0.01864 | 0.0082209 | 0.01864 | 0.0082209
| Day 4 | Run 1 | 91.6 | 398.7 | 91.6 | 398.7 | [14 empty] | SD B | 1.8919 | 8.3103 | 1.8919 | 8.3103
| Day 4 | Run 1 | 92.5 | 399 | 92.5 | 399 | [14 empty] | CV B | 0.020221 | 0.020287 | 0.020221 | 0.020287
| Day 4 | Run 1 | 91.2 | 398.2 | 91.2 | 398.2 | [14 empty] | SD WL | 2.5731 | 8.9667 | 2.5731 | 8.9667
| Day 4 | Run 1 | 93 | 402.4 | 93 | 402.4 | [14 empty] | CV WL | 0.027502 | 0.02189 | 0.027502 | 0.02189
| Day 5 | Run 1 | 92.9 | 401.4 | 92.9 | 401.4 | [14 empty] | DF WL | 9.3662 | 5.0658 | 9.3662 | 5.0658
| Day 5 | Run 1 | 92.1 | 402 | 92.1 | 402 | [14 empty] | ChiSq E | 36.7605 | 36.7605 | 36.7605 | 36.7605
| Day 5 | Run 1 | 94.7 | 404.2 | 94.7 | 404.2 | [14 empty] | ChiSq WL | 21.5956 | 14.658 | 21.5956 | 14.658
| Day 5 | Run 1 | 89.2 | 399.7 | 89.2 | 399.7 | [14 empty] | F Error | 1.3557 | 1.3557 | 1.3557 | 1.3557
| Day 5 | Run 1 | 88 | 398.6 | 88 | 398.6 | [14 empty] | F WL | 1.5185 | 1.701 | 1.5185 | 1.701
|  | Repeatability | 1.3206 | 1.8351 | 1.3206 | 1.8351
|  | Between day | 1.8919 | 8.3103 | 1.8919 | 8.3103
|  | Within lab | 2.5731 | 8.9667 | 2.5731 | 8.9667
|  | Repeatability CV | 1.864% | 0.8221% | 1.864% | 0.8221%
|  | Between day CV | 2.022% | 92.68% | 73.53% | 92.68%
|  | Within lab CV | 2.75% | 2.189% | 2.75% | 2.189%
Upper Verification Limits
|  | Repeatability UVL | 1.356% | 1.356% | 1.356% | 1.356%
|  | Within lab UVL | 3.948% | 4.423% | 3.948% | 4.423%
Comparison to Claim
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Repeatability | Fail | Pass | Fail | Pass
|  | Within Lab Precision | Pass | Pass | Pass | Pass
Assay Capability
|  |  | Level 1 | Level 2 | Level 3 | Level 4
|  | Mean | 2.5731 | 8.9667 | 2.5731 | 8.9667
|  | RCPA APS | #REF! | #REF! | #REF! | #REF!
|  | Capability (σ) | #REF! | #REF! | #REF! | #REF!
Acceptability
Comparison to Biological Variation APS
|  | CVi | #REF! |  |  | Specification
|  | CVg | #REF! |  | Imprecision
|  | Required Category | Desirable |  | Bias
|  | alpha | 0.05 |  | Total Error
Upper Verification Limits
|  |  | Level 1 | Level 2 | Level 3 | Level 4
Repeatability
Within Lab Precision
Imprecision Verification Result
Repeatability
Within Lab Precision
REFERENCE INTERVALS
Reference Intervals that will be used
Reference Interval Verification
| LRL | URL |  |  |  | LRL | URL
| 30 | 110 |  |  |  | 30 | 110
| Sample ID | Operator | Result |  |  | Sample ID | Operator | Result
1 | Test 1 | T1 | 61 | 0 |  | Test 1 | T1 | 61 | 0
2 | Test 2 | T2 | 97 | 0 |  | Test 2 | T2 | 97 | 0
3 | Test 3 | T3 | 93 | 0 |  | Test 3 | T3 | 93 | 0
4 | Test 4 | T4 | 48 | 0 |  | Test 4 | T4 | 48 | 0
5 | Test 5 | T5 | 52 | 0 |  | Test 5 | T5 | 52 | 0
6 | Test 6 | T6 | 52 | 0 |  | Test 6 | T6 | 52 | 0
7 | Test 7 | T7 | 41 | 0 |  | Test 7 | T7 | 41 | 0
8 | Test 8 | T8 | 75 | 0 |  | Test 8 | T8 | 75 | 0
9 | Test 9 | T9 | 123 | 1 |  | Test 9 | T9 | 123 | 1
10 | Test 10 | T10 | 55 | 0 |  | Test 10 | T10 | 55 | 0
11 | Test 11 | T11 | 56 | 0 |  | Test 11 | T11 | 56 | 0
12 | Test 12 | T12 | 111 | 1 |  | Test 12 | T12 | 111 | 1
13 | Test 13 | T13 | 74 | 0 |  | Test 13 | T13 | 74 | 0
14 | Test 14 | T14 | 73 | 0 |  | Test 14 | T14 | 73 | 0
15 | Test 15 | T15 | 67 | 0 |  | Test 15 | T15 | 67 | 0
16 | Test 16 | T16 | 95 | 0 |  | Test 16 | T16 | 95 | 0
17 | Test 17 | T17 | 55 | 0 |  | Test 17 | T17 | 55 | 0
18 | Test 18 | T18 | 44 | 0 |  | Test 18 | T18 | 44 | 0
19 | Test 19 | T19 | 59 | 0 |  | Test 19 | T19 | 59 | 0
20 | Test 20 | T20 | 112 | 1 |  | Test 20 | T20 | 112 | 1
21 | [5 empty] | Test 21 | T21 | 95 | 0
22 | [5 empty] | Test 22 | T22 | 90 | 0
23 | [5 empty] | Test 23 | T23 | 44 | 0
24 | [5 empty] | Test 24 | T24 | 65 | 0
25 | [5 empty] | Test 25 | T25 | 62 | 0
26 | [5 empty] | Test 26 | T26 | 46 | 0
27 | [5 empty] | Test 27 | T27 | 97 | 0
28 | [5 empty] | Test 28 | T28 | 56 | 0
29 | [5 empty] | Test 29 | T29 | 70 | 0
30 | [5 empty] | Test 30 | T30 | 71 | 0
31 | [5 empty] | Test 31 | T31 | 48 | 0
32 | [5 empty] | Test 32 | T32 | 64 | 0
33 | [5 empty] | Test 33 | T33 | 44 | 0
34 | [5 empty] | Test 34 | T34 | 75 | 0
35 | [5 empty] | Test 35 | T35 | 57 | 0
36 | [5 empty] | Test 36 | T36 | 52 | 0
37 | [5 empty] | Test 37 | T37 | 69 | 0
38 | [5 empty] | Test 38 | T38 | 95 | 0
39 | [5 empty] | Test 39 | T39 | 51 | 0
40 | [5 empty] | Test 40 | T40 | 74 | 0
|  | Reference interval verification | Fail |  |  |  | Reference interval verification | Pass
|  | N | 20 |  |  |  | N | 40
|  | Outliers | 3 |  |  |  | Outliers | 3
|  | p | 0.015902 |  |  |  | p | 0.13815
LIMIT OF BLANK, DETECTION, QUANTITATION
| Date of analysis |  |  | Reagent Lot No.
| Operator ID |  |  | Calibrator Lot No.
Limit of Blank (LoB)
| ENTER the LoB claim |  | 0.9 | #REF!
ENTER min of 2 replicates, of min 2 blank samples, of min 3 days, for min of 20 values
[9 empty] | Bounds for observed proportion of results relative to a detection capability claim | [10 empty] | Values ≤ LoB claim
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Total number of measurements (n) |  |  | % Observed proportion boundary | [7 empty] | 1 | 1 | 1
|  | Day 1 | 2022-10-23 | 0.5 | 0.8 |  |  |  | 20 |  |  | 85% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.5 | 0.5 |  |  |  | 30 |  |  | 86.67% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-23 | 0.6 | 0.6 |  |  |  | 40 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 2 | 2022-10-24 | 0.7 | 0.7 |  |  |  | 50 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.7 | 0.7 |  |  |  | 60 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-24 | 0.8 | 0.9 |  |  |  | 70 |  |  | 90% | [7 empty] | 1 | 1 | 1
|  | Day 3 | 2022-10-25 | 0.5 | 0.6 |  |  |  | 80 |  |  | 91.25% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.4 | 0.5 |  |  |  | 90 |  |  | 91.11% | [7 empty] | 1 | 1 | 1
|  |  | 2022-10-25 | 0.6 | 0.7 |  |  |  | 100 |  |  | 91% |  |  |  |  | n <= LoB claim |  | 27
[9 empty] | 150 |  |  | 92% |  |  |  |  | % n<= LoB claim |  | 150 | %
[9 empty] | 200 |  |  | 92.5%
| Total number of values |  | 18 |  | Min Req | 15 |  |  | 250 |  |  | 92.8%
| mean of blank |  | 0.62778 |  | % Min Req | 83.33% |  |  | 300 |  |  | 93%
| SD of blank |  | 0.12826 | [5 empty] | 400 |  |  | 93.25%
Limit of blank (LoB) |  |  | 0.84232 | 0 | Alpha | 0.05 |  |  | 500 |  |  | 93.4%
|  |  | Pass |  | C_beta | 1.6727 |  |  | 1000 |  |  | 93.8%
Limit of Detection (LoD)
| ENTER the LoD claim |  | 1.1 | #REF!
Values ≥ LoB claim
|  | ENTER min of 2 replicates, of min 2 LoD claim samples, of min 3 days, for min of 20 values | [17 empty] | 1 | 1
1
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  | Selected samples should have a concentration at the claimed LoD | [11 empty] | 1 | 1
|  | Day 1 | 2022-10-23 | 1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 0.9 | 0.8 | [14 empty] | 1 | 1
|  |  | 2022-10-23 | 1 | 1.1 | [14 empty] | 1 | 1
|  | Day 2 | 2022-10-24 | 1.1 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 0.9 | 1 | [14 empty] | 1 | 1
|  |  | 2022-10-24 | 1.1 | 1.1 | [14 empty] | 1 | 1
|  | Day 3 | 2022-10-25 | 1.2 | 1.1 | [11 empty] | n >=LoB claim |  | 17
|  |  | 2022-10-25 | 1.1 | 0.9 | [11 empty] | % n >= LoB claim |  | 94.4444 | %
|  |  | 2022-10-25 | 1 | 1.1
|  |  | Sample SD | 0.094281 | 0.099381
|  |  |  | 0.0088889 | 0.0098765 | #N/A
| Total number of values |  | 18 |  | Min Req | 15
| Pooled SD |  | 0.096864 |  | % Min Req | 83.33%
|  | Limit of Detection | 1.062 | 0
|  |  | Pass |  | Alpha | 0.05
[5 empty] | C_beta | 1.6727
Limit of Quantitaion (LoQ)
| Enter the LoQ claim |  | 1.5 | #REF!
ENTER min of 2 replicates, of min 2 LoQ claim samples, of min 3 days, for min of 20 values
% difference from means
|  | Day | Date | Sample 1 | Sample 2 | Sample 3 |  |  | Sample 1 | Sample 2 | Sample 3 |  |  |  |  | Values < TEa |  |  |  | % difference calculations
|  | Day 1 | 2022-10-23 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-23 | 13 | 14 |  |  |  | 14.5985 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 14.5985 | 8.0292 | #VALUE!
|  |  | 2022-10-23 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  | Day 2 | 2022-10-24 | 16 | 15 |  |  |  | 5.1095 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 5.1095 | 1.4599 | #VALUE!
|  |  | 2022-10-24 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-24 | 17 | 14 |  |  |  | 11.6788 | 8.0292 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 11.6788 | 8.0292 | #VALUE!
|  | Day 3 | 2022-10-25 | 15 | 16 |  |  |  | 1.4599 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 1.4599 | 5.1095 | #VALUE!
|  |  | 2022-10-25 | 14 | 15 |  |  |  | 8.0292 | 1.4599 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 8.0292 | 1.4599 | #VALUE!
|  |  | 2022-10-25 | 18 | 16 |  |  |  | 18.2482 | 5.1095 | #VALUE! |  |  |  |  | #REF! | #REF! | #VALUE! |  | 18.2482 | 5.1095 | #VALUE!
|  | Sample mean |  | 15.2222 | 15.2222 | [7 empty] | n >= TEa |  | 0
|  |  | Sample SD | 1.4741 | 0.78567 | [7 empty] | % n >= TEa |  | 0
|  |  | Sample CV | 9.684% | 5.161%
| Total number of values |  | 18 |  | Min Req | 15
| Total number of samples |  | 2 |  | % Min Req | 83.33%
Grand Mean |  |  | 15.2222
[5 empty] | Alpha | 0.05
| TEa |  | #REF! | % | C_beta | 1.6727
Fail
Note on B380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on C380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit
Note on G380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Lower reference limit
Note on H380: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Upper reference limit